Figure 2. Spectra of the signals processed inside the AIC structure. The AIC is able of reconstructing a signal even if its spectral content exceeds the first Nyquist region.

a)
b)

d)
c)

Figure 3. Estimated PSDs obtained by the FFT magnitude on a single record (a), or by averaging the FFT magnitudes evaluated on 10 (b), 100 (c), and 200 (d) records.
Figure 8. Block scheme of the AIC prototype and the test setup, used during experimental analysis.
Figure 6. Estimated PSDs of the reconstructed signal when infinite resolution is used (a), and when a quantization of 8 bit or 16 bit is used.
b)
a)

Figure 1: Block scheme of the random demodulation AIC.

a)
b)
c)
Figure 10. Evaluated figures of merit versus e and filter bandwidth, for a test signal frequency of 10 MHz.
Figure 9. Experimental results for test signal frequency ranging from 5 to 95 MHz: SFDR (a), SINAD (b), and THD (c), for e=0.5.
c)
b)
a)
Figure 11. PSD of the reconstructed signal for filter bandwidth equal to 14 and 20 MHz.
Figure 7. Estimated PSDs of the reconstructed signal for values of e equal to 0.01, 0.05, and 0.1. (a), respectively, and estimated PSDs of the reconstructed signal, for filter lengths equal to 17 and 49 taps (b).
a)
b)
Figure 5. Estimated PSDs of the reconstructed signal without ADC nonlinearity and with a nonlinearity coefficient c3 equal to 0.1, 0.2, and 0.3, respectively.
b)
a)

a)
b)

Figure 4. Estimated PSDs of the reconstructed signal without mixer nonlinearity (a), and with mixer nonlinearity (b).
# Frequency-domain characterization of random demodulation Analog-to-Information Converters
Doris Bao, Pasquale Daponte, Luca De Vito, Sergio Rapuano
Department of Egnineering, University of Sannio, Piazza Roma, 21, 82100 Benevento, Italy

The paper aims at proposing test methods for Analog-to-Information Converters (AICs).
In particular, the objective of this work is to verify if figures of merit and test methods, currently defined in standards for traditional Analog-to-Digital Converters, can be applied to AICs based on the random demodulation architecture.
For this purpose, an AIC prototype has been designed, starting from commercially available integrated circuits. A simulation analysis and an experimental investigation have been carried out to study the additional influencing factors such as the parameters of the reconstruction algorithm. Results show that standard figures of merit are in general capable of describing the performance of AICs, provided that they are slightly modified according to the proposals reported in the paper. In addition, test methods have to be modified in order to take into account the statistical behavior of AICs.

Keywords: Analog-to-Information Converter, Compressive sampling, Testing, Frequency domain.
Citation: Doris Bao, Pasquale Daponte, Luca De Vito, Sergio Rapuano, Frequency-domain characterization of random demodulation Analog-to-Information Converters, Acta IMEKO, vol. #, no. #, article #, Month Year, identifier: IMEKO-ACTA-##(Year)-##-##
Editor: Paolo Carbone, University of Perugia, Italy
Received month day, year; In final form month day, year; Published month year
Copyright: © 2012 IMEKO. This is an open-access article distributed under the terms of the Creative Commons Attribution 3.0 License, which permits unrestricted use, distribution, and reproduction in any medium, provided the original author and source are credited
Corresponding author: Luca De Vito, e-mail: devito@unisannio.it

Introduction
High-speed data acquisition is becoming a relevant topic in advanced applications, such as high-speed radar and communications, signal analysis, high-speed video acquisition, and so on. Moreover, it is a relevant challenge in wideband spectrum sensing for software defined radio and cognitive radio applications [1],[2].
Such demand often is not met by traditional Analog-to-Digital Converters (ADCs), due to technological limits in fast sampling rates [3].
The recent studies about compressive sampling (CS) drew a possible solution for signals, that can be represented by a finite number of non-zero elements in a specific domain.
They demonstrated that, for such class of signals, it is possible to reconstruct the original waveform, from a set of samples of a lower dimension than that required by the Shannon theorem.
The idea, underlying the AIC, is to spread the frequency content of the input signal. In this way, the high frequency components, folded back to low frequencies, can pass the anti-aliasing filter of the an ADC with a lower sampling frequency than that required by the Shannon's theorem for the original signal.
Basing on this concept, different architectures have been proposed, implementing the frequency spreading by exploiting: (i) non-uniform [4] or random sampling [5], (ii) random filters [6], and (iii) random demodulation [3].
The aim of the paper is to define performance parameters and test methods for AICs, starting from the state of art of research and the scientific knowledge about ADC testing, well summarized in [7] and [8]. To this aim, the first step is the application to AICs of standard parameters and test methods, actually defined for ADCs in order to study how they are influenced by (i) the AIC architecture type, (ii) the AIC design parameters, and (iii) the circuit non-idealities.
In the scientific literature few papers can be found, facing the AIC testing and most of them take into account only a reduced set of figures of merits and influencing parameters [9].
In [10], the authors presented a preliminary investigation carried out in simulation and on a first AIC prototype, based on a digital oscilloscope, about the application of standard ADC figures of merits on the AIC. This paper is an extended version of such work, in which new results are presented and a new AIC prototype is used, based on commercial integrated circuits.
As in [10], the AIC architecture considered in this paper is based on the random demodulation, as it does not require a high sampling frequency ADC. However, test methods and considerations can be easily extended to the other types of AIC architectures.
The paper is organized as follows: In Section 2, the theoretical background on compressive sampling is briefly recalled; in Section 3, the random demodulation architecture, which was used both in the simulation and the experimental analyses of this work, is described; in Section 4, the approach followed for AIC testing is explained; then, in Section 5, a simulation phase is reported, in which the influence of several factors, such as circuit non-idealities, AIC design parameters and reconstruction algorithm parameters, have been investigated. In Section 6, the figures of merit defined for traditional ADCs are revised. Finally, in Section 7, the experimental analysis is presented and results are discussed.
Theoretical background
The idea underlying CS approach is that many natural signals have concise representations when expressed in a convenient basis [11]. In many applications, in order to reduce the memory occupation or the transmission bandwidth, such signals are actually compressed, after being acquired using high speed or high resolution converters. The compression is carried out by transforming the signal to such domain where they have concise representation and discarding the coefficients close to zero.
CS aims to acquire directly the compressed version of the signal, without wasting acquisition or memory resources, by taking a vector y of observations of the signal to be acquired, where the size of y is lower than that required by the Shannon theorem to digitize the signal.
For a compressible signal x(t), if x is a the vector of N samples of it acquired according to the Shannon theorem, its compressed counterpart is represented by a vector y of size M < N, such that:
size 10 {bold nitalic {y} = bold nitalic {%PHI x},}	(1)
where, F is a matrix modelling the compression process.
It can demonstrated that x can be obtained from y according to (1) if a matrix transformation Y exists, such that:
size 10 {bold nitalic {x} = bold nitalic {%PSI c},}	(2)
where c has only K < M non-zero elements.
The above defined condition is not rare in reality, since many natural signals are sparse or compressible in the sense that they have concise representations when expressed in the proper basis.
By combining (1) and (2), the following expression is obtained:
size 10 {bold nitalic {y} = bold { nitalic {A c}},}	(3)
where A = F Y is an M × N matrix and, therefore, (3) is an under-determined linear system in c.
The system can be solved by finding the solution of (3) that minimizes the ℓ0 norm, that is having the highest number of non-zero elements in c:
size 10 {nitalic bold {hat c} = argmin lbrace {ldline nitalic bold {c} rdline}_{0} rbrace ~~~~{"subject to"}~~~~bold nitalic {y} = bold { nitalic {A c}}.}	(4)
The minimization of the ℓ0 norm is both numerically unstable and NP-complete, requiring an exhaustive enumeration of all size 10 {left ( binom{N}{K} right ) } possible locations of the non-zero entries in c [12]. Therefore, the solution is approximated with that obtained by the minimization of the ℓ1 norm:
size 10 {nitalic bold {hat c} = argmin lbrace {ldline nitalic bold {c} rdline}_{1} rbrace ~~~~{"subject to"}~~~~bold nitalic {y} = bold { nitalic {A c}}.}	(5)
In presence of additive noise e, equation (3) becomes:
size 10 {bold nitalic {y} = bold {nitalic {A c}} + nitalic bold {e}}	(6)
and the minimization problem is modified as:
size 10 {nitalic bold {hat c} = argmin lbrace {ldline nitalic bold {c} rdline}_{1} rbrace ~~~~{"subject to"}~~~~ ldline {bold nitalic {y} - bold { nitalic {A c}}} rdline _2 < %tau .}	(7)
or as:
size 10 {nitalic bold {hat c} = argmin lbrace {ldline nitalic bold {c} rdline}_{1} rbrace ~~~~{"subject to"}~~~~ ldline {bold nitalic {A^T} left ( bold nitalic {y} - bold { nitalic {A c}} right ) } rdline _infty < %varepsilon .}	(8)
where t and e are small positive constants.
The ℓ1 norm minimization is a convex optimization problem that conveniently can be reduced to a linear program known as basis pursuit, whose computational complexity is about O(N 3) [12].
The minimum number of observations M required to successfully reconstruct x is given by the following expression [11]:
size 10 {M geslant C %mu^2 (bold{%PHI}, bold{%PSI}) K log N,}	(9)
where, C is a positive constant and size 10 {%mu (bold{%PHI}, bold{%PSI})} is a quantity, called coherence between F and Y, having the following expression:
size 10 { %mu (bold{%PHI}, bold{%PSI}) = sqrt N max csub{1 leslant i,j leslant N} lline langle %phi_i, %psi_j rangle rline, }	(10)
with ji and yj the i-th row of F and the j-th column of Y, respectively, and size 10{langle {} cdot {} ,{} cdot {} rangle} the dot product.
By looking at (9), it can be easily seen that the lower is the coherence between F and Y, the fewer observations are needed to reconstruct the original signal. It can be demonstrated that incoherence can be achieved with high probability simply by selecting F as a random matrix [12], whose elements are drawn from a suitable distribution, such as Gaussian, Bernoulli or Rademacher.
The random demodulation AIC architecture
The block scheme of the random demodulation AIC is shown in Fig.1. Considering an ADC with sampling frequency fs, and a signal x(t), whose frequency content exceeds the first Nyquist region size 10 {left [0,f_s over 2 right ],} it would not be possible to sample x(t), according to the Shannon theorem (Fig. 2a). It can be observed that the condition (2) is verified as soon as x has a Fourier representation with a limited number of non-zero coefficients. In that case, in fact (2) is simply the Fourier synthesis formula.
Therefore, first, a compressed version of x is could be obtained, and, then, the signal could be reconstructed by solving one of the problems (5), (7) or (8).
In order to obtain the compressed version of x(t), it is mixed with a Pseudo-Random Binary Sequence (PRBS), having a bit rate equal to fp = ℳ fs, with ℳ integer, The spectrum of the PRBS, shown in Fig. 2b, has a frequency content spread over the whole region size 10 {left [0,f_p right],} therefore, the output of the mixer contains the information of the original signal, spread over the whole first Nyquist band, even if it was originally located outside such region (Fig. 2c). This allows that after low-pass filtering (Fig. 2d) and A/D conversion (Fig. 2e), the information contained in the original signal has not been lost. However, it needs to be recovered from the filter output.
A reconstruction algorithm is, then, in charge of extracting this information, by solving one of the problems in (5), (7) or (8), as previously quoted (Fig. 2f).
A mathematical model of the AIC can be obtained as it follows. Indicating with x the vector of the signal samples, if it was acquired using a sampling frequency equal to the bit rate of the PRBS fp, and with size 10{ nitalic bold p = left [ matrix {p_0, # p_1, # dotslow, #p_{N-1}} right ] } the samples of the PRBS, the output of an ideal digital mixer m can be obtained by multiplying x by a diagonal matrix, having the PRBS samples on the diagonal:
size 10 {nitalic bold {m} = nitalic bold {D x}}	(11)
where:
size 10 {nitalic bold {D} = left [ matrix{p_0 # 0 # dotsaxis # 0 ## 0 # p_1 # dotsaxis # 0 ## dotsvert # dotsvert # dotsdown # dotsvert ## 0 # 0 # dotsaxis # p_{N-1}} right ] }	(12)
Instead, the filter can be modelled by the following Toeplitz matrix:
size 10 {nitalic bold {H} = left [ matrix{ h_1 # h_2 # dotsaxis # h_{L-1} # 0 # dotsaxis # 0 ## 0 # h_1 # dotsaxis # h_{L-2} # h_{L-1} # dotsaxis # 0 ## dotsvert # dotsvert # dotsdown # dotsvert # dotsvert # dotsaxis # dotsvert ## 0 # 0 # dotsaxis # dotsaxis # dotsaxis # h_{L-2} # h_{L-1} } right ] }	(13)
where, size 10 {nitalic bold {h} = left [ matrix{h_0 # h_1 # dotsaxis # h_{L-1}} right ]^T } is the truncated impulse response of the antialiasing filter of length L. L should be chosen large enough that the truncated response guarantees a good approximation of the actual one.
The random demodulation AIC can be obtained by choosing a F matrix that decimates the matrix product H D, by taking only 1 every ℳ rows.
Approach to AIC Testing
AIC testing is a harder task than traditional ADC testing, due to the following reasons: (i) the AIC output signal has a strong stochastic behaviour, due to the PRBS (compare Figs.2a and 2e); (ii) the AIC output signal is noise-like for every input, thus, it is hard to analyse either in the time or in the frequency domain; and (iii) the parameters of the reconstruction algorithm can affect the AIC results.
The current research has been focused on frequency-domain testing, as it is mostly used in telecommunications, which is the main application field of AICs. The approach followed for AIC testing consists of using a sinewave signal as input to the AIC, applying a reconstruction algorithm, and characterizing the spectrum of the reconstructed signal, by means of the traditional ADC figures of merit. Test signal frequencies have been chosen such that the coherent sampling condition is respected for the reconstructed signal, that is they should verify the following condition:
size 10{ %omega_0 = {J over N} 2 %pi, }	(14)
where, J is an integer number, relatively prime to N.
The ADC testing methods in the frequency domain can be clearly applied to AIC as the involved signals are sparse in such domain. As sparsity in the frequency domain has been considered, the Y matrix in (2) must represent a transformation between time domain and frequency domain. However, since only real-valued test signals have been considered, the Discrete Hartley Transform (DHT) basis has been used, which is equivalent to the Discrete Fourier Transform, but is purely real.
The generic element yi,j of the DHT has the following form:
size 10{ %psi_{i,j}={1 over sqrt N }left [ cos left ( {2%pi}over N ij right ) + sin left ( {2%pi}over N ij right ) right ] }.	(15)
The Dantzig selector [13] has been used as reconstruction algorithm, solving the optimization problem in (8), by means of the implementation given by the ℓ1 -magic Matlab toolbox [14].
Differently from problem (7), which looks for the solution in the least square sense, the Dantzig selector computes the residual r = y – A c, and, then, evaluates its inner products with the columns of A. These inner products form various weighted combinations of the entries in r and all those are expected to be small [15]. The constraint, thus, proceeds by requiring that all these are below the threshold e.
Therefore, the parameter e represents the maximum allowable value for the correlation of the residual r with any column of A.
In a preliminary phase, the AIC has been modelled and some simulations have been carried out in order to understand how the figures of merit, evaluated on the reconstructed signal, can be affected by: (i) circuit non-idealities, (ii) test signal parameters, and (iii) reconstruction algorithm parameters.
According to the results of the simulation analysis, the following figures of merit, have been extended to AICs: Spurious Free Dynamic Range (SFDR), Signal to Noise and Distortion Ratio (SFDR), and Total Harmonic Distortion (THD).
For such figures of merit, the definitions in [16] have been used. However, the same approach can be used even considering definitions coming from different standards, such as those considered in [17].
Finally, an experimental phase has been conducted by evaluating the revised figures of merit on a specifically designed AIC prototype.
SIMULATION analysis
During simulation analysis, an explicit behavioural model [18] has been considered for each AIC block. As any actual mixer will distort the signals, the ideal mixer output yid = x1(t) x2(t) (see Eq. 11) has been replaced by a second order polynomial function, as follows:
size 10{ y(t) = a x_1 (t) + b x_2 (t) + x_1 (t) x_2 (t) + c x_1^2 (t) + d x_2^2 (t) }	(16)
where, x1(t) and x2(t) are the inputs to the mixer, respectively, and a, b, c and d are attenuation coefficients.
The antialiasing filter has been modelled as a lowpass FIR filter, with cut-off frequency equal to p/4.
Finally, the ADC has been modelled as a memoryless nonlinearity followed by an ideal quantizer, where the nonlinearity has been obtained, as in [9], using the following third order polynomial function:
size 10 { y(t)={x(t)+c_3 x^3 (t)} over {1+c_3} }	(17)
Simulation tests have been performed by acquiring 200 records, each with N=512 samples, using a ratio between the PRBS bit rate and the ADC sampling frequency equal to 8.
An estimate of the output signal Power Spectral Density (PSD), obtained by averaging the Fast Fourier Transforms (FFTs) of the 200 acquired records has been used to estimate the frequency domain parameters.
  1. Influence of average number for PSD estimation
In Fig. 3, the estimated PSD is shown, when it is obtained by the FFT (Fast Fourier Transform) magnitude of a single record (Fig. 3a), or when it is obtained by averaging FFT magnitudes evaluated on different records. Results are shown for a number of averages equal to 10 (Fig. 3b), 100 (Fig. 3c), and 200 (Fig. 3d). As it can be noted, due to the selection of frequency components operated in the reconstruction algorithm, some of them are not identified in a single record. Instead, the averaging allows obtaining an estimation of the PSD for each frequency component.
Such behaviour is more evident for higher values of e, since in this case a higher number of components are discarded by the reconstruction algorithm.
  1. Influence of mixer nonlinearity
The main effect of the mixer nonlinearity is of increasing the level of the noise floor, as shown in Fig.4, where the PSDs of the reconstructed signal when no mixer nonlinearity has been modelled (a=b=c=d=0) (Fig.4a), and when a nonlinearity is added, with the following values of the attenuation coefficients: a=b=0.1, c=d=0.01.
Those values have been chosen by looking at the order of magnitude of the typical feedthrough and second order distortion component attenuation for actual mixers.
In addition, a particular increase of the level of the components, whose frequencies lie at multiples of the fundamental, has been observed.
  1. Influence of ADC nonlinearity
The effect of ADC nonlinearity can be observed in Fig. 5, where the PSDs of the reconstructed signals are reported in the case the c3 parameter is set to 0 (Fig.5a), simulating a linear ADC, and when the c3 parameter is set to 0.2 (Fig. 5b). Differently from traditional ADCs, the nonlinearity does not act only on a restricted number of components, located at multiples of the fundamental, in fact it causes even an increase of the noise floor.
  1. Influence of ADC quantization
The ADC quantization causes an increase of the noise floor level. As it can be seen in Fig. 6a, the signal is perfectly reconstructed (with noise floor level depending only on the machine precision) when maximum machine resolution is used. Instead, a noise floor arises when quantization is used. However, only a small difference in the level has been observed for different resolutions. During simulation analysis, values of the resolution in the set {8, 10, 12, 14, and 16} bits have been used. In Fig. 6b, the PSDs, related to a resolution of 8 bits and 16 bits are shown.
  1. Influence of reconstruction parameter e
As the reconstruction parameter e represents a limit to the amount of additive noise that can be tolerated on the output signal, it controls the selection of the frequency components in the reconstructed signal. In fact, for each record, a frequency component will be included in the reconstructed signal, only if its energy exceeds a certain level, depending on the value of e.
On the averaged spectrum, this causes an increase of the level of both the noise floor and of the harmonics, for decreasing values of e.
Such behaviour can be observed in Fig. 7a, where the PSDs of the reconstructed signal are shown, where e has been set to 0.01, 0.05, and 0.1.
  1. Influence of the filter length
The length of the impulse response of the anti-aliasing filter determines the capability of the filter of attenuating the aliasing replicas. The shorter the filter impulse response is, the lower such capability is. Fig.7b highlights this behaviour, by showing the PSDs of the reconstructed signals, when filter lengths of 17 and 49 taps are used. As it can be noted, in the case of a filter length of 17 (blue line), the filter is not able to adequately attenuate the aliasing replicas, located at the frequencies size 10{ left lbrace {2 %pi} over %em - %omega_0, {2 %pi} over %em + %omega_0 right rbrace . } In the case the filter length is of 49 taps (red line), instead, the aliasing components are well attenuated under the noise floor.
Definition of the figures of merit
As previously mentioned, the figures of merit used in this work have been defined starting from those of the IEEE Std. 1241-2010 [16]. However, the following modifications have been introduced: (i) the number of averages of the spectral magnitudes has been set to 200, in order to take into account the stronger stochastic behaviour of the reconstructed signal, and (ii) the frequencies size 10{ left lbrace {2 %pi} over %em - %omega_0, {2 %pi} over %em + %omega_0 right rbrace } have been included in the list of the harmonics, contributing to the THD.
The figures of merit have been evaluated by averaging the spectral magnitude size 10{left lline X[m] right lline} as it follows:
size 10 { X_avm[m] = { 1 over R sum from {k=1} to R {left lline X_k [m] right lline} }, ~~ m=0, 1, 2, dotslow, N-1, }	(18)
over a number of R = 200 data records.
Compared to definitions in [16], no changes have been made to the SFDR and SINAD:
size 10 { SFDR = 20 log_10 { left ( {left lline X_avm[n_i] right lline} over {max csub {m in S_0} {left lline X_avm[m] right lline}} right ) } }	(19)
size 10 { SINAD = 20 log_10 { sqrt { { {N-3} over N } cdot { {X_avm^2[n_i] + X_avm^2[N-n_i]} over {sum csub {m in S_0} {X_avm^2[m]}} } } } }	(20)
where, ni is the frequency index corresponding to the fundamental component, S0 is the set of frequency indexes from 1 to N -1, excluding the two values corresponding to the fundamental and its image.
Instead, THD has been modified as follows:
size 8.7 { THD = 20 log_10 { sqrt { { sum from {h=2} to N_h {X_avm^2 left [n_h right ]} + sum from {h=-N_h} to -2 {X_avm^2 left [n_h right ]} + X_avm^2 left [ left lline M-n_i right rline right ] + X_avm^2 left [ mod(M+n_i,2) right ] + X_avm^2 left [ N- left lline M-n_i right rline right ] + X_avm^2 left [ N-mod(M+n_i,2) right ] } over {X_avm^2 left [ n_i right ] + X_avm^2 left [ M-n_i right ]} } } }	(21)

where, Nh is the number of the highest harmonics considered, mod(·,·) is the modulo operator, and nh is the frequency index corresponding to a generic harmonic or its aliasing component:
size 10 { n_h = mod(h n_i,N)~~h=+-[2, 3, dotslow, N_h] }	(22)
In the considered case Nh has been chosen equal to 10.
Experimental investigation
For the purpose of the experimental investigation, an AIC prototype has been designed, using commercially available circuits and bench instrumentation. In the following Subsections, the AIC prototype is described, then, some details, about the calibration phase needed to estimate the F matrix, are given; finally, the experimental results are presented.
  1. The AIC prototype
A block scheme of the AIC prototype is shown in Fig. 8. The mixing process is carried out by means of the Analog Devices AD8342 evaluation board, while the antialiasing filter has been realized by the first channel of the Analog Devices ADRF6510 evaluation board. The ADRF6510 is a dual-channel programmable lowpass filter, with cut-off frequency tunable in the range [1-30] MHz. After mixing and filtering, signal is acquired by means of an Analog Devices AD9230 evaluation board, operating at a sampling frequency of 50 MSa/s. The ADC board is connected with a computer through an Analog Devices HSC-ADC-EVALC Capture board.
For the test purpose, sinewave signals have been generated by an Agilent E4438C Vector Signal Generator, while the PRBS is generated in the computer and downloaded, through the GPIB, to the Tektronix AWG420 Arbitrary Waveform Generator. An Agilent E8663B is used to generate the sampling clock for the ADC board.
In order to synchronize the PRBS generation with the acquisition, the HSC-ADC-EVALC Capture board has been configured with a specific firmware allowing external triggering. In this case, the acquisition is started on the computer, causing a Ready signal to be set on the HSC-ADC-EVALC Capture board. Such Ready signal is given both to the Tektronix AWG420, to generate the PRBS and again to the HSC-ADC-EVAL Capture board to start the acquisition. The Agilent E4438C, the Tektronix AWG420 and the Agilent E8663B are then synchronized by means of a 10 MHz reference, in order to guarantee the coherent sampling condition, and controlled by the computer by means of a GPIB interface.
  1. F matrix estimation
In order to run the reconstruction algorithm, it is necessary to estimate the F matrix in (1). Since it is obtained from the PRBS and from the impulse response of the anti-aliasing filter through (12) and (13), only the estimation of the impulse response of the anti-aliasing filter is necessary.
This is done in a preliminary calibration phase, by connecting the Tektronix AWG420, generating a PRBS, directly to the input of the filter. Then, the output of the filter is acquired. As shown in [19], if the input to a system x(t) is a stochastic process, having impulsive autocorrelation function, the impulse response of the system can be measured as:
size 10 {h(n) = {r_xy(n)} over {r_xx(0)},}	(23)
where, rxy(n) is the cross-correlation function between the input and the output and rxx(0) is the autocorrelation of the input, evaluated at a lag equal to 0, corresponding to the input signal energy. Since in the considered case the generated PRBS signal has impulsive autocorrelation, it is possible to estimate the impulse response of the filter using (23).
  1. Experimental results
Using the AIC prototype and the test setup shown in Fig. 8, a wide experimental investigation has been carried out, by evaluating the SFDR, SINAD, and THD, on the reconstructed signal. As in simulation tests, sinewave input signals have been used. In order to guarantee the coherent sampling condition, the test signal frequency has been selected, using the following expression:
size 10 {f_0 = {J over N} cdot {f_p} }	(24)
where, fp is the PRBS bit rate, corresponding to the equivalent sampling frequency, and J is an integer coprime to N.
Some results are shown in Fig. 9, for a test signal frequency in the range [5‑95] MHz, with a step of 5 MHz.
The variations of the values of all three figures of merit are in the range of some dBs, even beyond the first Nyquist band [0‑25] MHz, however, performance show a little decrease, in SFDR and SINAD, for frequencies approaching 100 MHz.
It is worth studying the variation of the figures of merit versus the reconstruction parameter e, as well as versus the bandwidth of the anti-aliasing filter.
As shown by the simulation results, the levels of both the noise floor and the harmonic components decrease with e, therefore, an increase in the values of SINAD and SFDR is expected, while THD is expected to decrease. This can be easily observed in Fig. 10, where the trend of the considered figures of merit versus e is shown.
In Fig. 10, the variation of the figures of merit versus the bandwidth of the anti-aliasing filter is shown, too. It can be observed that the values of SINAD and THD are generally better for a filter with a wider band. This is probably due to the fact that a greater amount of information passes through the filter and can be used for the reconstruction. In the case of the SFDR, instead, smaller values have been observed with the narrowband filter, and e < 1. As it can be seen in Fig. 11, this due to the aliasing components at fs - f0 and fs + f0, that are not attenuated under the noise floor level, by the 20 MHz bandwidth filter. In the case of a 20 MHz bandwidth, in fact, the noise floor level is lower, but the aliasing component are higher than the case of a 14 MHz bandwidth.
Conclusions and further work
In this paper, a frequency-domain characterization of AICs has been presented, by concentrating the attention on the random demodulation architecture. In a former phase, the AIC has been simulated by means of explicit behavioural models of its main components, in order to evaluate the influence of the circuit non-idealities, the design parameters and the reconstruction algorithm to the figures of merit, currently defined for ADCs.
The simulation results show that both mixer and ADC nonlinearities contribute to increase the noise floor level, as well as the level of the harmonic components. This is quite different from what happens in traditional ADCs, where nonlinearities do not affect the noise floor level. In addition, the antialiasing filter can affect the SFDR, as it can not completely attenuate the aliasing components. As result of the simulation analysis, a modification to the THD definition is proposed by including the energy of the aliasing components.
In a latter phase, an experimental analysis has been conducted on a specifically designed AIC prototype, using sinewaves, by varying the test input frequency, the reconstruction parameter e, and the bandwidth of the antialiasing filter. Such analysis showed that the proposed figures of merit are in general capable of describing the performance of the AIC. However, the figures of merit should be observed for different values of e, since high values of e can hide the effects of noise and distortion.
Further work will be directed to analyse the Intermodulation Distortion as frequency domain figures of merit. Then, it can be interesting to observe the figures of merit even evaluated in the time domain. In addition, the methodology presented in this paper can be extended to other AIC architectures, such as those based on the random sampling. Finally, an alternative approach to AIC testing can be used, without relying on the reconstruction algorithm, but directly comparing the AIC output, with the expected output of the AIC model.
References
1. L.De Vito, “Methods and technologies for wideband spectrum sensing”, Measurement, vol.46, No.9, Nov.2013, pp.3153–3165.
2. D.Bao, L.De Vito, S.Rapuano, “A histogram-based segmentation method for wideband spectrum sensing in cognitive radios”, IEEE Trans. on Instrum. and Meas., vol. 62, No.7, July 2013, pp.1900–1908.
3. S.Kirolos, J.Laska, M.Wakin, M.Duarte, D.Baron, T.Ragheb, Y.Massoud, R.Baraniuk, “Analog-to-information conversion via random demodulation”, Proc. of 2006 IEEE Dallas/CAS Workshop on Design, Applications, Integration and Software, Richardson, US, Oct. 2006, pp.71–74.
4. M.Wakin, S.Becker, E.Nakamura, M.Grant, E.Sovero, D.Ching, J.Yoo, J.Romberg, A.Emami-Neyestanak, E.Candes, “A nonuniform sampler for wideband spectrally-sparse environments”, IEEE Journal on Emerging and Selected Topics in Circuits and Systems, vol.2, No.3, Sept. 2012, pp. 516–529.
5. J.Laska, S.Kirolos, “Random sampling for analog-to-information conversion of wideband signals”, Proc. of IEEE Dallas/CAS Workshop on Design, Applications, Integration and Software, Richardson, US, Oct. 2006, pp.119–122.
6. J.A.Tropp, M.B.Wakin, M.F.Duarte, D.Baron, R.G.Baraniuk, “Random filters for compressive sampling and reconstruction”, Proc. of IEEE Int. Conf. on Acoustics Speed and Signal Processing, Toulouse, France, May 2006, vol.3, pp.III–872–III–875.
7. S.Rapuano, P.Daponte, E.Balestrieri, L.De Vito, S.J.Tilden, S.Max, J.Blair, “ADC parameters and characteristics - Part 6 in a series of tutorials in instrumentation and measurement”, IEEE Instrum. Meas. Magazine, vol.8, No.5, 2005, pp. 44–54.
8. T.E.Linnenbrink, J.Blair, S.Rapuano, P.Daponte, E.Balestrieri, L.De Vito, S.Max, S.J.Tilden, “ADC testing - Part 7 in a series of tutorials in instrumentation and measurements”, IEEE Instrum. Meas. Magazine, vol.9, No.2, 2006, pp. 39–49.
9. Z.Yu, J.Zhou, M.Ramirez, S.Hoyos, B.M.Sadler, “The impact of ADC nonlinearity in a mixed-signal compressive sensing system for frequency-domain sparse signals”, Physical Communication, vol.5, No.2, June 2012, pp.196–207.
10. D.Bao, P.Daponte, L.De Vito, S.Rapuano, “Defining frequency domain performance of Analog-to-Information converters”, Proc. of 17th IMEKO Int. Workshop on ADC Modelling and Testing, Barcelona, Spain, July 2013, pp.748-753.
11. E.J.Candes, M.B.Wakin, “An introduction to compressive sampling”, IEEE Signal Processing Magazine, vol.25, March 2008, pp.21–30.
12. R.G.Baraniuk, “Compressive Sensing [Lecture Notes]”, IEEE Signal Processing Magazine, vol.24, N.4, July 2007, pp.118–121.
13. E.Candes, T.Tao, “The Dantzig selector: Statistical estimation when p is much larger than n”, The Annals of Statistics, vol.35, No.6, Dec.2007, pp.2313–2351.
14. ℓ1 -magic toolbox, http://users.ece.gatech.edu/~justin/l1magic/.
15. M.Elad, “Sparse and redundant representations: from theory to applications in signal and image processing”, Springer, New York, 2010.
16. “IEEE Standard for terminology and test methods for analog-to-digital converters”, IEEE Std. 1241, 2011.
17. S.Rapuano, “Preliminary considerations on ADC standard harmonization”, IEEE Trans. on Instrum. and Meas., vol.57, No.2, Feb.2008, pp.386-394.
18. P.Arpaia, P.Daponte, S.Rapuano, “A state of the art on ADC modelling”, Computer Standards and Interfaces, vol.26, No.1, Jan.2004, pp.31-42.
19. K.J.Keesman, “System identification: an introduction”, Springer-Verlag, London, 2011.