Research on volume determination of mass standards with two acoustic measuring chambers

**Man Hong Hu, Jian Wang, Yue Zhang, Chang Qing Cai, Rui Lin Zhong, Jing An Ding, and Xiao Ping Ren**
Division of Mechanics and Acoustic, National Institute of Metrology, No.18, Bei San Huan Dong Lu, Chaoyang Dist, Beijing, P.R.China

**• The major points of extension in a supplementary file:**

In the extended paper, the major points of extension are shown in Figure 7 and Table 5. Comparing with the paper in APMF 2015, relative high acoustic volume measuring accuracy has been experimentally achieved by reducing the residual air in the acoustic chambers.

|  |
| --- |
| (a) |
|  |  |
| (b) | (c) |
| Figure 7 After reducing the remaining air in the chambers (a), The 1 kg reference weight and test weights ranging from 100 g to 5 kg (b) and the measuring results (c)  |

Table 5. Uncertainty budget of test weights using 1 kg weight as reference weight after reducing the air in the acoustic chambers

|  |  |
| --- | --- |
| Sources  |  Test weights |
| 5 kg  | 2 kg  | 1 kg  | 500 g  | 200 g  | 100 g |
| Volume of reference weight(cm3)  | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 |
| *R*1(cm3)  | 0.068 | 0.028 | 0.017 | 0.009  | 0.005 | 0.002 |
| *R*2(cm3)  | 0.014 | 0.008 | 0.005  | 0.004 | 0.002  | 0.001 |
| *R*3(cm3)  | 0.084 | 0.041 | 0.026  | 0.012  | 0.006 | 0.003 |
| Combined uncertainty(cm3)  | 0.109 | 0.051 | 0.031 | 0.016 | 0.008 | 0.005 |
| Extended uncertainty(cm3)  | 0.218 | 0.102 | 0.062 | 0.033 | 0.017 | 0.010 |
| Relative extended uncertainty( *k*=2, ×10-2)  | 0.035 | 0.041 | 0.049 | 0.052 | 0.068 | 0.078 |