Research on volume determination of mass standards with two acoustic measuring chambers

Man Hong Hu, Jian Wang, Yue Zhang, Chang Qing Cai, Rui Lin Zhong, Jing An Ding, and Xiao Ping Ren 
Division of Mechanics and Acoustic, National Institute of Metrology, No.18, Bei San Huan Dong Lu, Chaoyang Dist, Beijing, P.R.China

• The major points of extension in a supplementary file:

In the extended paper, the major points of extension are shown in Figure 7 and Table 5. Comparing with the paper in APMF 2015, relative high acoustic volume measuring accuracy has been experimentally achieved by reducing the residual air in the acoustic chambers.

| (a) |  |
| --- | --- |
| (b) | (c) |
| Figure 7 After reducing the remaining air in the chambers (a), The 1 kg reference weight and test weights ranging from 100 g to 5 kg (b) and the measuring results (c) |  |

Table 5. Uncertainty budget of test weights using 1 kg weight as reference weight after reducing the air in the acoustic chambers

| Sources | Test weights |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 5 kg | 2 kg | 1 kg | 500 g | 200 g | 100 g |
| Volume of reference weight(cm3) | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 | 0.003 |
| R1(cm3) | 0.068 | 0.028 | 0.017 | 0.009 | 0.005 | 0.002 |
| R2(cm3) | 0.014 | 0.008 | 0.005 | 0.004 | 0.002 | 0.001 |
| R3(cm3) | 0.084 | 0.041 | 0.026 | 0.012 | 0.006 | 0.003 |
| Combined uncertainty(cm3) | 0.109 | 0.051 | 0.031 | 0.016 | 0.008 | 0.005 |
| Extended uncertainty(cm3) | 0.218 | 0.102 | 0.062 | 0.033 | 0.017 | 0.010 |
| Relative extended uncertainty( k=2, ×10-2) | 0.035 | 0.041 | 0.049 | 0.052 | 0.068 | 0.078 |