
30 May to 1 June, 2017, Helsinki, Finland
Proposal and investigation of a new method for calibration of strain cylinders using laser interferometry 
Junning Cui1, Rolf Kumme2*/Presenter, Holger Kahmann2
1 Harbin Institute of Technology, Harbin, PR China, cuijunning@hit.edu.cn

2 Physikalisch-Technische Bundesanstalt, Braunschweig, Germany, rolf.kumme@ptb.de
Abstract: This research is to investigate the possibility of calibrating a strain cylinders using laser interferometry, thus to provide a new type of transducer that can provide both force and deformation indications. This new method for calibration of strain cylinders is based on application of a double channel laser interferometer in a force standard machine. With the new method proposed, the relationship between force, deformation and strain can be calibrated in parallel when calibration forces are applied according to procedure used in ISO 376. Experimental results show that the deformation of a strain cylinder has a definite and stable relationship with the force applied, and can be calibrated and directly traced to the wavelength of laser. So a new standard which can be used for both alignment verification and indication verification of compression testing machines and other uniaxial testing machines is proved. This research also illustrates the possibility of providing a new deformation-type force transducer using a non-gauged steel cylinder together with a multi-channel laser interferometer.
Keywords: strain cylinder; laser interferometry; calibration; deformation; compression testing machine.
1. INTRODUCTION
A strain cylinder, also named strain gauged column, is a kind of force transducer with four channel outputs. It is often used for alignment verification of compression testing machines. Specifications of strain cylinders and proving procedure for verification of compression testing machines are specified in international and national standards [1-2]. The calibration of stain cylinders can be done in a force standard machine according to international standard ISO 376 [3].
Stain cylinders are usually used in two ways. First, since the difference between signals of four bridges represent the distribution of force applied on strain cylinders, they are used as standards for alignment verification of compression testing machines, including self-alignment of upper machine platens, alignment of component parts of machines, and so on. Second, they can also be used as normal force transducers for force measurement taking the average signals of four bridges. So during indication calibration of compression testing machines, stain cylinders can also be used as force transfer standards. Normally, only mechanical qualities like force and stain are interested in during calibration and application of strain cylinders, while deformation is not paid special attention to. However, since the platens of testing machines are of big size and strong enough, integral deformation is ignorable, and compression deformation of strain cylinders and local deformation distribution in the surface of platens is mostly decided by material characteristics. Therefore, in many applications the deformation is of great significance and it is often hoped deformation indication is provided.
Methods and tools in precision engineering are frequently introduced into the field of force metrology In recent years. For example, laser interferometry has been used in measurement of deformation or displacement of key parts [4, 5], dynamic calibration [6-8], and so on. Coordinate measuring technique has been used in on-machine or off-machine measurement of key dimensional parameters [9, 10]. These methods and tools show great capability and advantages in solving some certain problems. In this paper, a new method for calibration of strain cylinders using laser interferometry is proposed and investigated.

In this paper, a new method for calibration of strain cylinders using laser interferometry is proposed and investigated. The new method is based on the application of a double channel laser interferometer in a force standard machine, and shows advantages and provides more information. 
2.  Basic concepT
The structure of a stain cylinder and strain bridges on it is shown in Fig. 1. A strain cylinder consists of a steel cylinder, 16 strain gauges and other accessories. High demands are required for its material, manufacturing, assembling and calibration. The steel cylinder should be precisely manufactured and assembled with 16 matched and temperature compensated electrical resistance strain gauges. Each 4 strain gauges are wired in a form of a full bridge, electrically and thermally balanced, and assembled at one of the ends of a pair of orthogonal diameters half-way up the cylinder. Furthermore, the strain cylinder needs to be precisely calibrated on a force standard machine together with the dedicated strain measuring equipment. This calibration qualifies the strain cylinder to be used as a standard for both alignment verification and indication verification of testing machines.
[image: image1.png]



Fig. 1 Structure of a stain cylinder and strain bridges on it [1]
During calibration and application of stain cylinders, conversions of physical quantities are shown in Fig. 2. It can be seen that a number of conversions of physical quantities are involved, during which deformation and strain serve as the most important intermediate quantities. At present, stain cylinders are calibrated and used with interest only focused on the strain rather than deformation. However, in many applications the deformation is of great significance and represent important information about alignment, force distribution, and so on. That the reason why it is often hoped that deformation indication is also provided by stain cylinders. Normally, platens of compressing testing machines are designed in such a way that it is big and strong enough to avoid integral deformation, so local deformation in the surface of platens and strain cylinders is only related to their material characteristics. Therefore, important information can be derived from data of deformation.
[image: image2.png]Force standard/
testing machine

Force

Steel

Deformation

Strain

Resistance

strain meas.

Voltage in mV/y

>

leylinder]

and strain

gauges

equipment

— Conversion of value





Fig. 2 Conversions of physical quantities of stain cylinders
The idea of the new method for calibration of strain cylinders is to introduce laser interferometer in a force standard machine and obtain deformation information when calibration forces are applied, and calibrate the relationship between force, deformation and strain in parallel. In this way, a new standard that can provide both force and deformation indication is provided, and can be used for both alignment verification and indication verification of compression testing machines is proved. Also this research illustrates the possibility of providing a new deformation-type force transducer using a non-gauged steel cylinder together with a multi-channel laser interferometer. This new force transducer is much simpler in structure and usage not requiring the complicated manufacturing and assembling process of strain gauges.
This idea is applicable on the condition that the deformation of a strain cylinder has a definite and stable relationship with the force applied, and can be calibrated using a laser interferometer in a force standard machine with good repeatability and reproducibility. This will be proved through experiments and assessment of this new type of transducer in following sections.
3.  Experimental setup and calibration procedure
3.1 Experimental setup
The experimental setup is shown in Fig. 3. The experiments are carried out in a 5 MN force standard machine. The maximum force applied on the cylinder is 2 MN. The strain cylinder, laser heads and reflectors of the laser interferometer are assembled between the upper platen and lower platen of the machine, and the strain cylinder is centrally assembled on the lower platen.
[image: image3.png]Low platen Strain cylinder Laser heads

a)

Strain cylinder

Bridge

Meas. points
of laser heads

Low platen

b)




a)

[image: image4.png]Low platen Strain cylinder Laser heads

a)

Strain cylinder

Bridge

Meas. points
of laser heads

Low platen

b)




b)
Fig. 3 Experimental setup in the 5MN force standard machine. a) Experimental setup. b) Top view of strain cylinder with 4 bridges and 4 corresponding measurement points 
The strain cylinder is 100 mm in diameter and 200 mm in height. It conforms to requirements about material and dimensional tolerances in Ref. [1]. Four full bridges of strain gauges, numbered as Channel 1 to Channel 4, are assembled at fours ends of a pair of orthogonal diameters half-way up the cylinder. Each bridge consists of two elements measuring the axial strain and two elements measuring the circumferential strain. The outputs of four bridges are recorded by dedicated strain measuring equipment DMP41.
A MI-5000 double channel laser interferometer from SIOS Meßtechnik GmbH is used. The relative length measurement accuracy of it is up to 1(10-6, so the length measurement error introduced is negligible during experiments, showing great advantage of a tool from the field of precision engineering. For each channel, interferometric optics excluding the measuring reflector are assembled together as a laser head. Laser is transmitted to each of the two laser heads through a length of optical fiber. Corresponding to orientations of four bridges on the strain cylinder, there are four corresponding measurement points for laser heads on the lower platen, which are at fours ends of a pair of corresponding orthogonal diameters of a circle and designated as “A”, “B”, “C” and “D” in Fig. 3 b). In such a way, the strain cylinder together with the laser heads work as a new type of transducer with four “deformation sensing elements”. The four “deformation sensing elements” are uniformly distributed over 360( and numbered as Channel 1 to Channel 4 too.
The distance between measurement points and the sidewall of the strain cylinder should be kept as small as possible. In this research it is 60 mm. The influence of this distance and resulted deformation distribution in the surface of platens can be taken into consideration in two ways: 1) when the strain cylinder is used in a compression testing machine, keep the same parameter. Usually, the platens of the compression testing machine and the force standard machine are made of material with similar Young’s Elastic Modulus. Correct the error introduced if there is obvious difference; 2) correct the deformation to the center position of platens for both the force standard machine and the compression testing machine.
3.2 Calibration procedure
Procedure of calibrating a strain cylinder using a double channel laser interferometer is proposed with reference to ISO 376. The cylinder is rotated around its axis to four uniformly distributed positions over 360(, as shown in Fig. 4. Deformation and strain data at the four rotation positions are averaged to eliminate the influence of non-uniform force distribution characteristics of the force standard machine. A series of increasing calibration forces and then a series of decreasing calibration forces are applied at each rotation position. Because stain cylinders are only used for compression test, data corresponding to deceasing calibration forces are not used for interpolation, but only used for assessment of relative reversibility error in Section 4.3. The two laser heads are first assembled at Point “A” and Point “B”, and the strain cylinder is calibrated at four rotation positions. Then the two laser heads are re-assembled to Point “C” and Point “D”, and the strain cylinder is calibrated at another four rotation positions. In such a way, outputs of four channel “deformation sensing elements” of the new type of transducer can be obtained at each rotation position.
The series of increasing calibration forces are from 0 to 2000 kN with 200 kN increment, and the series of decreasing calibration forces are from 2000 kN down to 0 with 200 kN decrement. Totally 11 calibration forces including zero force which are uniformly distributed over the calibration range are applied. There is a wait of ~ 15 s after a calibration force is applied, waiting for the force to become stable. Then the calibration force is kept unchanged for ~ 45 s before it is changed to next successive value.
Data recorded by DMP41 are designated as Xchi, j and Xchi, j( in the following sections, where i is the channel number of strain bridges and i = 1, 2, 3 or 4, j is the rotation position of the strain cylinder and j = 0, 90, 180 or 240, Xchi, j is corresponding to increasing calibration forces and Xchi, j( is corresponding to decreasing calibration forces. Similarly, data recorded by the laser interferometer are designated as Ychi, j and Ychi, j(, where i is the channel number of “deformation sensing elements”.
[image: image5.png]C TN N

Cne
[en cial @B

7 cn3 s <
ez o) @B » [os el @B » {ore o) @B »
N \cny / \ N/ O

u

0° rotation position 90° rotation position 180° rotation position 270° rotation position | Reassembling
of laser heads

Ce Ce Co Co

.- e

NN

« {os o b * (o ot « Cfom co) b
\ 3 \ Chd, S

N \ch2 NN .





Fig. 4 Procedure of calibrating a strain cylinder using a double channel laser interferometer
3.3 Data sampling setting
It is a good idea to use high data sampling rate based on the strategy of “more data, more information”. However, higher sampling rate bring forward higher demand for hardware and software, and make the data processing more difficult. On the other hands, according to knowledge about the force standard machine, the natural frequency of micro vibration of the lower platen is ~ 0.27 Hz, so the period of vibration is ~ 3.7 s. Therefore, the data sampling rate of the laser interferometer and DMP41 is set to be 10 Hz. So the sampling can be regarded as “dynamic” sampling compared to the period of platen vibration, without omitting any important information.
4.  Data processing and experimental results
4.1 Pre-processing of  data
The abundance of data continuously sampled at a sampling rate of 10Hz are pre-processed first, so to obtain values of deformation and strain corresponding to 11 calibration forces during each incremental force loading and decremental force loading.
For data of laser interferometer during each calibration, first, corner points corresponding to the moments when a calibration force begins to increase or decrease to the next successive value are determined with a slope-judging algorithm, marked with green “*” in Fig. 5. Then the average values of 25 s data before each corner point are calculated, marked with green lines in Fig. 5. This averaging operation works as a low pass filter (LPF) with a cutoff frequency of ~ 0.04 Hz, thus to eliminate the influence of noise resulted from mechanical vibration and other factors. Deformation values corresponding to 10 calibration forces during each incremental force loading and decremental force loading are obtained.
Theoretically, the indication of the laser interferometer should be reset to zero at the moment when the upper platen comes into contact with the strain cylinder, i.e. the moment when the calibration force begins to increase from zero. It is a good solution to send a trigger signal to the laser interferometer either automatically or manually, or following two method can be used: (1) during data pre-processing, recognize the data corresponding to zero force and subtract all successive data by it, which is the case in this research; (2) use values of deformation corresponding to 10 calibration forces from 200 kN to 2000 kN for interpolation, then calculate deformation value from 0 to 200 kN through extrapolation, and subtract all deformation values from 0 to 2000 kN by the value corresponding to zero force.
[image: image6.png]Laserfum

500

400

300

—s— Channel 2
—s— Channel 4

Averaging line
of 255y

Corner points.-

1000 1500
Time/second

2000 2500




Fig. 5 Pre-processing of data of laser interferometer (0( rotation position)
The data of DMP41 are pre-processing similarly, as shown in Fig. 6. The only exception is that the strain value corresponding to zero force can be directly calculated through averaging of 25 s data before the corresponding corner point. 11 strain values corresponding to 11 calibration forces including zero force are used for interpolation.
[image: image7.png]DMP41/(mVA)

15

05

Comner points |

500

1000 1500
Time/second

2000

2500




Fig. 6 Pre-processing of four channel data of DMP41 (0( rotation position)
4.2 Data processing results
Through data pre-processing, four channel values of deformation and strain corresponding to 11 calibration forces at each rotation position, either increasing or decreasing, are obtained. The interpolation curves are determined from average values of deformation and strain at four rotation positions. As mentioned before, data corresponding to decreasing calibration forces not used for interpolation. Three-degree polynomial fitting are carried out to obtain third degree equations about relationship between force, deformation and strain.
4.2.1 Polynomial fitting results of force vs. deformation
Equations giving the deformation as a function of strain are calculated as follows:

[image: image8.wmf]632

111

632

222

632

333

632

44

-2.013*10*0.002064*4.272*-0.4283

-1.748*10*0.001859*4.337*0.3774

-2.263*10*0.002216*4.270*0.06657

-2.356*10*0.002293*4.226*

 

 

chchch

chchch

chchch

chch

FYYY

FYYY

FYYY

FYY

-

-

-

-

=++

=+++

=+++

=++

4

-0.7880

ch

Y

ì

ï

ï

í

ï

ï

î

(1)

[image: image9.wmf]632

-2.096*10*0.0021083*4.276*-0.2008

FYYY

-

=++

  (2)
where F is the calibration force in unit of kN, Ychi is the deformation of Channel i in unit of (m, i = 1, 2, 3 or 4, 
[image: image10.wmf]Y

is the average deformation of the four channels.
Curves of the four three-degree equations are shown in Fig. 7. It can be seen that the 4 channels of the strain cylinder have very good consistency.
[image: image11.png]—— Channel 1
2000 | —=— Channel 2
—— Channel 3

Z —*— Channel 4
1500 r —— Average
o
2
£ 1000
1220
500 1200
1180 =555 60—
ot

0 50 100 150 200 250 300 350 400
Deformation / #m




Fig. 7 Curves of relationship of force vs. deformation
It can be concluded that the deformation of a strain cylinder has a definite and stable relationship with the force applied by the force standard machine. Equations (1) and (2) also illustrate the possibility of providing a new deformation-type force transducer using a non-gauged steel cylinder together with a multi-channel laser interferometer, using either average or non-averaged values of deformation of four channels. The average value represents the amplitude of the force applied on the strain cylinder, while non-averaged values of four channels illustrate the force distribution.
4.2.2 Polynomial fitting results of deformation vs. strain and force vs. strain
Equations giving the deformation as a function of strain are calculated as follows:

[image: image12.wmf]32

1111

32

2222

32

3333

32

4444

8.584*-31.77*283.3*0.2128

3.058*-15.29*268.0*-0.5468

7.588*-28.62*279.4*0.03244

13.35*-45.43294.3*1.062

chchchch

chchchch

chchchch

chchchch

YXXX

YXXX

YXXX

YXXX

ì

=++

ï

=+

ï

í

=++

ï

ï

=++

î

    (3)

[image: image13.wmf]32

8.102*-30.22281.3*0.1592

YXXX

=++

              (4)
where Xchi is the strain indication of Channel i in unit of mV/V, i = 1, 2, 3 or 4, 
[image: image14.wmf]X

is the average strain of the four channels.
Curves of the four three-degree equations are shown in Fig. 8. Since the four strain bridges are well matched and thermally and electrically balanced, so these four curves have good consistency too.
[image: image15.png]—— Channel 1
400 —+— Channel 2
—*— Channel 3
g —*— Channel 4
= 300 —*— Average
£
3 I
g 200 280 T
S 260 /
2100 240
L 095 1 1.05
0 1 1 1 1 1 1 1 1 1
0 0.4 0.8 1.2 1.6

Strain / (mV/V)




Fig. 8 Curves of relationship of deformation vs. strain
Equations giving the calibration force as a function of stain can be similarly calculated as follows:

[image: image16.wmf]32

111

32

222

32

333

32

444

3.942*-9.654*1221*0.04842

-15.24*48.81*1168*-2.339

-3.916*12.59*1204*-0.2183

19.09*-55.75*1261*3.191

chchch

chchch

chchch

chchch

FXXX

FXXX

FXXX

FXXX

ì

=++

ï

=++

ï

í

=++

ï

ï

=++

î

   (5)

[image: image17.wmf]32

0.8159*-0.7826*1214*0.06329

FXXX

=++

     (6)
During real applications, either average signal or non-averaged signals of four channel strain bridges can be used. The average signal represents the amplitude of the force applied on the strain cylinder, while the non-averaged values of four channels illustrates the force distribution applied on the strain cylinder. 
With Equations (3) ~ (6), a new type of transducer which provides both force indication and deformation indications is provided. This new type of transducer can be used as a standard for both alignment verification and indication verification of compression testing machines. 
4.3 Assessment of deformation calibration results
The calibration results of deformation Y are assessed with reference to ISO 376. The relative length measurement accuracy of the MI-5000 laser interferometer is up to 1(10-6, the uncertainty of calibration forces applied by the 5MN force standard machine is up to 1(10-4 (k = 2), so this assessment can be used to verify the new method proposed. Data corresponding to decremental force loading are used during assessment of relative reversibility error.
4.3.1 Relative reproducibility error
Relative reproducibility error of deformation value Y assesses the variation of deformation indication at different rotation positions. Averaged deformation of the four channels is used for assessment. The error is calculated using the following equations in analogy to ISO 376:

[image: image18.wmf]ï

ï

î

ï

ï

í

ì

+

+

+

=

´

-

=

4

100

min

max

270

180

90

0

Y

Y

Y

Y

Y

Y

Y

Y

b

ave

ave

                       (7)
The relative reproducibility errors are shown in Fig. 9. The maximum relative reproducibility error bmax happens when the calibration force is 200 kN, bmax = 6.4. It has been found that the reproducibility error is mostly resulted by the non-uniform force distribution characteristics of the force standard machine. The indication variation of deformation at different rotation positions and the relative reproducibility error can be significantly reduced if the force distribution is more uniform. 
[image: image19.png]800 1200 1600 2000

400

L L L L L L o
~ O v T n a4 — o

9% / 10119 Ayijiqronpoxdar oAne[oy

Force / kN




Fig. 9 Relative reproducibility errors in %

4.3.2 Relative interpolation error
Relative interpolation error is calculated for each channel using the following equation:

[image: image20.wmf]100

´

-

=

chi

chi

chi

chi

Ya

Ya

Y

fc

                           (8)
where 
[image: image21.wmf]chi

Ya

 is the value of deformation of Channel i computed through inverse functions of Equations 1. These functions give the deformation as a function of calibration force. 
The relative interpolation errors of Channel 1 to Channel 4 are shown in Fig. 10. The maximum relative interpolation error |fc|max happens in Channel 4 when the calibration force is 200 kN, |fc|max = 0.84. The relative interpolation error can be averaged if the averaged deformation of the four channels is used.
[image: image22.png]Relative reproducibility error / %

0.8
0.6
04
02t

-0.2 1

04

—— Channel 1
—+— Channel 2 1
—*— Channel 3

*— Channel 4

400 800 1200 1600 2000
Force / kN




Fig. 10 Relative interpolation errors of 4 channels
4.3.3 Relative zero error
The relative zero error of Channel i at j degree rotation position is calculated using the following equation:

[image: image23.wmf]100

,

,

,

,

´

-

=

j

chi

j

chi

j

chi

j

chi

YN

Yo

Yf

fo

                       (9)
where 
[image: image24.wmf]j

chi

Yo

,

 is the deformation taken before the calibration force is applied, 
[image: image25.wmf]j

chi

Yf

,

 is the deformation taken after the calibration force is completely removed. 
[image: image26.wmf]j

chi

YN

,

 is the deformation corresponding to the 2000 kN maximum calibration force. 
The relative zero errors of Channel 1 to Channel 4 are shown in Fig. 11. The maximum relative zero error fomax happens in Channel 3 when the cylinder is at 270( rotation position, fomax = -3.5
[image: image27.png]Relative zero error / %

Ky

41

21

———

———

Channel 1
Channel 2
Channel 3
Channel 4

90 180
Rotation position / Deg

270





Fig. 11 Relative zero errors of 4 channels
4.3.4 Relative reversibility error
The relative reversibility error of Channel i at j degree rotation position 
[image: image28.wmf]j

chi

v

,

 and the relative reversibility error of Channel i 
[image: image29.wmf]chi

v

 are respectively calculated using the following equations:


[image: image30.wmf],,

,

,

,0,90,180,270

100

4

chijchij

chij

chij

chichichichi

chi

YY

v

Y

vvvv

v

°°°°

¢

ì

-

ï

=´

ï

í

ï

+++

ï

=

î

        
 MACROBUTTON MTPlaceRef \* MERGEFORMAT (1)

The relative reversibility errors of Channel 1 to Channel 4 are shown in Fig. 12. The maximum relative reversibility error vmax happens in Channel 4 when the calibration force is 200 kN, vmax = 13. The reversibility error is mostly resulted by the hysteresis between incremental and decremental force loading. Since strain cylinders are only used for compression test, this does not proved to be a problem. 
[image: image31.png]Relative reproducibility error / %

14
—— Channel 1
12r —<— Channel 2
—— Channel 3
10 —*— Channel 4 A
gl J
6 J
4t J
o J
0 400 800 1200 1600 2000

Force / kN




Fig. 12 Relative reversibility errors of 4 channels
4.3.5 Relative creep error
The relative creep error of Channel i is calculated using the following equation:

[image: image32.wmf]100

20

200

´

-

=

chi

chi

chi

chi

YN

Y

Y

c

                      (11)
where 
[image: image33.wmf]chi

Y

20

 and 
[image: image34.wmf]chi

Y

200

 are deformation indications obtained 20 s and 200 s after the maximum calibration force is applied, 
[image: image35.wmf]chi

YN

 is the maximum deformation corresponding to the 2000 kN maximum calibration force. Remark: According ISO 376 the creep should be measured for 300 s but in this experiment to reduce measurement time the creep was measured within 200 s.
Average calculation of 25 s data which works as a 0.04 Hz low pass filter is used. The experimental results of relative creep error test are shown in Fig. 13. The maximum relative creep error cmax = 0.091. 
[image: image36.png]Laser / (4m

460

440

420 ¢

400

380 1

360

/'

The moment of
force application of 25s

Averaging line

/'

340 —e—Channel 1 1
120 . . ‘ ‘ —O—Cflanvnel 3
0 50 100 150 200 250 300

Time / second




Fig. 13 Experimental results of relative creep error test
4.4 Proposals
In this research, calibration procedure, data processing method and assessment of transducer calibration is designed with reference to ISO 376. However, some suggestions and proposals should be paid attention to:
(1) The cylinder should be rotated symmetrically round its axis to at least three uniformly distributed positions over 360(, thus to eliminate the influence of non-uniform force distribution characteristics of the force standard machine;
(2) In this research, incremental and decremental force loading are both applied on the strain cylinder. Since strain cylinders are only used for compression test, data corresponding to decremental force loading are only used for assessment rather than interpolation. In order to get as much information as possible, “dynamic” data sampling is used, this is not mandatorily needed;
(3)
During data pre-processing, calculation of deformation and strain corresponding to zero force should be paid special attention to. Readings of strain bridges can be directly taken before the force is applied or after the force has been completely removed. However, readings of the laser interferometer corresponding to zero force should be taken at the moment that the upper platen comes into contact or loose contact with the cylinder. The indication of the laser interferometer can be reset to zero at this moment can be done using the methods recommended in Section 4.1;
(4) If a non-gauged steel cylinder is calibrated with the method proposed, and to be used together with a multi-channel laser interferometer thus to provide a “deformation-type” force transducer. The distance between measurement laser beams and the sidewall of the cylinder should be kept as small as possible, and the error introduced can be considered using the methods recommended in Section 3.1. If the deformation of platens is negligible because the calibration force is too small or the platens are strong enough, this does not prove to be a problem. 
5.  Conclusion
A new method for calibration of strain cylinders using laser interferometry is proposed and investigated. This new method is illustrated through calibrating a strain cylinder with a SIOS double channel laser interferometer in a 5 MN force standard machine at PTB. Calibrated procedure, data processing method and assessment of calibration results are designed with reference to ISO 376. Experimental results show that the deformation of a strain cylinder has a definite and stable relationship with the force applied, and can be calibrated and directly traced to the wavelength of laser.
Therefore, a new type of transducer that can provide both force and deformation indications is proved, and can be used for both alignment verification and indication verification of compression testing machines and other uniaxial testing machines. This method can also be applied to non-gauged standard cylinders, thus to provide a new deformation-type force transducer using a non-gauged steel cylinder together with a multi-channel laser interferometer.
6. Acknowledgement
The authors gratefully acknowledge the valuable help of Heiko Wunderlich of PTB on experimental operation, and constructive technical discussion with Michael Wagner of PTB. 
7. Reference
[1]
EN 12390-4: Testing hardened concrete – Part 4: Compressive strength – Specification for testing machines, 2000.

[2]
DIN 51302-2: Strain cylinder test for verifying compression testing machines for testing concrete, 2000.
[3]
ISO 376: Metallic materials – Calibration of force-proving instruments used for the verification of uniaxial testing machines, 2011.

[4]
Gordon A. Shaw, Koo-Hyun Chung, Douglas T. Smith etc. Methods for transferring the SI unit of force from millinewtons to piconewtons. Proceedings of the SEM Annual Conference, Albuquerque New Mexico USA. 2009.
[5]
C. Bartoli,_, M.F. Beug, T. Bruns etc. Traceable dynamic measurement of mechanical quantities: objectives and first results of this european project. Int. J. Metrol. Qual. Eng. 3, 127–135 (2012).

[6]
Li Zhang, Rolf Kumme. Investigation of interferometric methods for dynamic force measurement. Proceedings, XVII IMEKO, Dubrovnik, Croatia. 2003: 315-318.

[7]
Yusaku Fujii, J D R Valera. Impact force measurement using an inertial mass and a digitizer. Meas. Sci. Technol. 17 (2006) 863–868.

[8]
Yusaku Fujii. Optical method for accurate force measurement: dynamic response evaluation of an impact hammer. Optical Engineering 45(2): 023002-1 ~ 023002-7.

[9]
Sebastian Baumgarten, Dirk Röske, Holger Kahmann etc. Concept and setup of a multi-component facility for force and torque in the range of 1 MN and 2 kN·m.  ACTA IMEKO. 2(3), 2014: 39-43.

[10]
D. Röske, K. Adolf and D. Peschel. Lever optimization for torque standard machines. XVI IMEKOm, Vienna, AUSTRIA, 2000, September 25-28.
7
823Y-2

_1555433574.unknown

_1555433578.unknown

_1555433582.unknown

_1555433585.unknown

_1555433587.unknown

_1555444403.unknown

_1555433586.unknown

_1555433584.unknown

_1555433580.unknown

_1555433581.unknown

_1555433579.unknown

_1555433576.unknown

_1555433577.unknown

_1555433575.unknown

_1543824424.unknown

_1555433568.unknown

_1555433571.unknown

_1543824438.unknown

_1543824455.unknown

_1543824296.unknown

_1543824416.unknown

_1543824292.unknown

