Validation of a PTB force-balanced piston gauge primary pressure standard

Ahmed S. Hashad1, Wladimir Sabuga2, Sven Ehlers2, Thomas Bock3

1 National Institute for Standards (NIS), Giza, Egypt (Guest researcher at PTB Germany)  
2 Physikalisch-Technische Bundesanstalt (PTB), Braunschweig, Germany  
3 Physikalisch-Technische Bundesanstalt (PTB), Berlin, Germany

ABSTRACT

Experimental methods using different pressure standards were applied to verify theoretical results obtained for the effective area of a piston-cylinder assembly (PCA) and for pressures measured with a force-balanced piston gauge (FPG). The theoretical effective area was based on the PCA’s dimensional properties defined by the diameter, straightness and roundness measurements of the piston and cylinder, derived from gas-flow modelling, using principles of rarefied gas dynamics and presented as two values, one for absolute and the other for gauge-pressure operation mode. Both values have a relative standard uncertainty of 5⋅10-6. The experimental methods chosen were designed to cover the entire operating pressure range of the FPG from 3 Pa to 15 kPa. Comparisons of the FPG with three different PTB pressure standards operated in different pressure ranges – a pressure balance, a mercury manometer and a static expansion system – were performed using the cross-float method and by a direct comparison of the generated pressures. For the theoretical and experimental effective area, as well as for pressures generated by the FPG and the reference standards, all the results demonstrated full agreement within the expanded uncertainties of the standards.

Section: RESEARCH PAPER

**Keywords:** FPG; cross floating; effective area; pressure standard; experimental verification

**Citation:** Ahmed S. Hashad, Wladimir Sabuga, Sven Ehlers, Thomas Bock, Validation of a PTB force-balanced piston gauge primary pressure standard, Acta IMEKO, vol. A, no. B, article C, Month Year, identifier: IMEKO-ACTA-A (Year)-B-C

**Section Editor:** Momoko Kojima, NMIJ, Japan

**Received** month day, year; **In final form** month day, year; **Published** Month Year

**Copyright:** This is an open-access article distributed under the terms of the Creative Commons Attribution 3.0 License, which permits unrestricted use, distribution, and reproduction in any medium, provided the original author and source are credited.

**Funding:** This work was supported by the EMPIR-participating countries within EURAMET and the European Union and by the Egyptian study mission.

**Corresponding author:** Paul P. L. Regtien, e-mail: paul@regtien.net

1. Introduction


Force-balanced piston gauges (FPGs) present an important class of piston gauge that enables pressure measurement from a few pascals to several tens of kilopascals. In these instruments, the force of the pressures acting on the piston of a piston-cylinder assembly (PCA) is measured by a high-accuracy balance. This force and the effective area () of the PCA can be used to calculate the difference in the pressures above and below the piston [1], [2]. At PTB, a force-balanced piston gauge manufactured by Fluke Calibration (USA), model FPG8601 [3], [4], was characterised as a primary pressure standard by determining of the PCA from the PCA’s dimensional properties [5], [6]. For this, a rarefied gas dynamics model was used to determine the pressure distribution of the gas flow inside the gap between the piston and the cylinder [7]-[9]. Three experiments were then performed to validate the PCA’s effective area and the FPG pressures by comparing them with the pressures of three different standards. These standards have various operation principles and pressure ranges, as elaborated on below.

1. Experiments

Experiments that can be used for the metrological characterisation of an FPG are described in, for example, [3], [10]-[17]. In the present study, experiments were specified that allowed the verification of the metrological characteristics obtained by the theoretical calculations. In the following sections, these experiments are described in detail. In all the experiments, a very low-pressure controller (VLPC), produced by Fluke Calibration, was used to stabilise the FPG measurement and lubrication pressures, unless otherwise stated.

* 1. Effective area determination

The FPG was calibrated against a Ruska pressure balance, model 2465A (gas-piston gauge), equipped with a PCA, serial no. TL1568, to determine the effective area of the FPG by the cross-float method, as described in [18], [19]. The Ruska PCA has an effective area equal to 3.35666 × (1 ± 5.8⋅10‑6) cm2 (all uncertainties in this paper are given for  = 1, unless otherwise specified). The effective area of this Ruska PCA is traceable to the state-of-the-art pressure balance that was used in the experiments on the redetermination of the Boltzmann constant [20]-[22] and is the German national pressure standard in the 180 kPa to 7.5 MPa range of absolute pressure. Both instruments, Ruska and FPG, were connected via the setup shown in Figure 1.

(a) (b)

Figure 1. Experimental setup for cross floating between FPG and Ruska pressure balance: sketch of the whole setup (a) and life photo of the experiment (b).

A class E2 stainless-steel mass set was used to characterise the FPG balance [23]. A 10 torr capacitance diaphragm gauge (CDG) was placed in the measurement line to indicate the pressure difference between the two instruments. The pressures of both instruments were controlled to maintain the pressure difference measured with the CDG as close to zero as possible. The measurements were performed in gauge and absolute mode in the pressure ranges of 3 kPa to 15 kPa and 2 kPa to 15 kPa, respectively. The higher starting pressure point of 3 kPa in gauge mode was chosen due to the Ruska pressure balance’s poor performance at lower pressures.

During the measurement, valve V1, which separated both instruments, was closed, and valve V2 was opened. The pressure of the FPG was generated by the VLPC at the target pressure, and the variable pressure volume was used to generate the target pressure within the Ruska pressure balance. The VLPC then finely adjusted the pressure within the FPG until the CDG indication was close to zero. Finally, valve V1 was opened and valve V2 closed to allow a direct cross float between the FPG and the Ruska pressure balance.

From the experimental results, the effective area was calculated by means of equation (1),

|  |  |
| --- | --- |
|  | (1) |

in which is the pressure of the reference standard, is the residual pressure on the lower side of the FPG PCA, which is equal to zero in gauge mode, is the thermal expansion coefficient of the FPG PCA and is the force measured by the balance and corrected for buoyancy and lubrication pressure-change effects, as set out in [24] and shown in equation (2),

|  |  |
| --- | --- |
|  | (2) |

where is the true mass of the internal reference mass used for the balance’s internal calibration, is the reading of the balance when is loaded, is the indication of the balance resulting from the pressure difference, is the indication correction due to the lubricant gas-pressure variation, is the indication correction due to the drag-force change and is the indication correction due to the atmospheric-pressure variation.

The Ruska reference pressure () was calculated by means of equation (3),


|  |  |
| --- | --- |
|  | (3) |

where is the local gravity acceleration, and are the masses and their densities, respectively, is the density of the ambient gas, is the piston’s additional volume (which is required as a correction due to the buoyancy produced by the pressure-transmitting medium) with density , is the effective area of the Ruska PCA, is the thermal expansion coefficient of the Ruska PCA, is the residual pressure in the Ruska bell jar and is the height difference between the reference levels of the pressure balance and the FPG.

* 1. FPG vs mercury manometer

In this experiment, the FPG was calibrated against the PTB mercury manometer (HgM), model 1025B made by Schwien Engineering (USA), which was modified at PTB and metrologically characterised, as described and validated in [25]‑[27]. The design of this instrument is shown in Figure 2. This manometer comprises two vessels filled with mercury and connected by a flexible tube (A) with the ability to change the height of one of the vessels mechanically. The change in height of the movable vessel is measured by means of a laser interferometer (B). The height levels of the mercury’s free surface in the vessels are controlled by capacitance measurements (C1 and C2). In the experiments, the FPG was connected to the HgM directly; this is demonstrated in Figure 3 for measurements in gauge mode.

A CDG was placed between the two instruments to measure the pressure difference when changing the pressure in each instrument. All connections were thermally isolated to minimise the effect of ambient temperature changes even though all measurements were performed in an air-conditioned environment with a temperature stability within 0.3 K. The measurements were carried out at a pressure range from 100 Pa to 15 kPa in the gauge mode and from 1 kPa to 15 kPa in the absolute mode.

(a) (b)

Figure 3. Experimental setup for cross floating between the FPG and the HgM in gauge-pressure mode: sketch of the whole setup (a) and life photo of the experiment (b).



Figure 2. PTB mercury manometer (HgM) with its reference (-) and measurement (+) pressure lines and climatisation temperature sensors (x).

In the gauge-pressure mode, the VLPC was used as usual. At the beginning, zero pressures were set in both instruments, and both instruments were zeroed. After zeroing, pressures were generated separately. When generating the pressures, valve V1, which separated both instruments, remained closed, and valve V2 was open, allowing the HgM to be connected to its pressure generator. Thus, the target pressure was generated by the VLPC within the FPG and by the pressure generator within the HgM. The height of the movable vessel was adjusted to make capacitances C1 and C2 equal (see Figure 2). Valve V2 was then closed and, using the VLPC to finely adjust the pressure in the FPG, a close-to- zero indication of the CDG was reached. Next, bypath valve V1 was opened to establish a direct connection and a pressure equilibrium between both systems. When equilibrium was achieved and all relevant indications became stable, their values and all conditions were recorded.



Figure 4 Experimental setup for cross floating between FPG and HgM in absolute mode.

In the absolute-pressure mode, the VLPC was replaced with valve V4, which has a finely controlled opening and closing mechanism, behind which a turbopump was located (see Figure 4). The lubrication pressure was equal to 40 kPa, allowing the measurement pressure inside the FPG to be calculated by means of equation (4),

|  |  |
| --- | --- |
|  | (4) |

in which is the height difference between the mercury levels in the two vessels, is the height difference between the reference level of the HgM and the mercury surface in the lower vessel, is the density of mercury, is the density of the pressure-transmitting gas and is the residual pressure in the upper vessel of the HgM. The FPG pressure was calculated by means of equation (5) with two additional components, containing the height difference between the pressure reference levels of the FPG and HgM and the residual pressure on the reference side of the FPG, which appears only in the absolute-pressure mode,

|  |  |
| --- | --- |
|  | (5) |

* 1. FPG vs static expansion system

The static expansion system (SES) [28], [29], see Figure 5, is a primary vacuum standard consisting of two stages of expansion, starting with two small volumes (*VL*1 and *VL*2) and two intermediate volumes (*VL*4 and *VL*5) and concluding with a calibration vessel (*VL*6) with a nominal volume of 100 litres. UUC is the unit under calibration. This set of volumes enables the realisation of the expansion ratios , which are presented in Table 1.

Table 1. Expansion ratios of the SES shown in Figure 5.

|  |  |  |
| --- | --- | --- |
| **Symbol** | **Expansion ratio** | **f** |
| f1 |  | 9.1732⋅10-3 |
| f2 |  | 7.4231⋅10-4 |
| f3 |  | 9.1978⋅10-3 |
| f4 |  | 7.4378⋅10-4 |
| f5 |  | 9.170⋅10-3 |

A Quartz Bourdon Spiral (QBS) was used to control and measure the initial pressure value when filling the first volume with the nitrogen gas in a range between 1 kPa and 100 kPa. The final pressures produced by SES in the range of 10 mPa and 1 kPa can be calculated by means of equation (6),

|  |  |
| --- | --- |
|  | (6) |

where is the corrected expansion ratio, taking the additional volume of the UUC into account, is the initial temperature before and the final temperature after the expansion,  is the virial coefficient and is the gas constant.

The SES was used to verify the FPG in the low-pressure range from 3 Pa up to 300 Pa. Because of the humidity of the nitrogen used inside the FPG and in order to keep the expansion volume *VL*6 defined, a CDG, with a full range of 13.33 kPa and 1 mPa resolution, constantly separated the two systems, and its readings were taken into consideration in the comparison results.




Figure 5. Design of SES.

(a) (b)

Figure 6. Experimental setup for comparison between FPG and SES: sketch of the whole setup (a) and life photo of the experiment (b).

Figure 6 shows the experimental setup for the comparison of the FPG with the SES in absolute-pressure mode, where the SES is sketched simply with only two volumes. The VLPC was used to keep the pressure difference measured by the CDG as close to zero as possible. Initially, the QBS was used to set the pressure in *VLS*1. Valves V2 and V6 were then opened to expand the gas into the larger volume *VLS*2 and the tubes connected to the left side of the CDG before V1. The total volume was measured precisely. The SES pressure value was calculated according to equation (6).

1. Results

The results of the effective area determination based on the FPG calibration against the Ruska pressure balance, which was described in Section 2.1, are presented in Table 2 and Figure 7 for both gauge and absolute-pressure operation modes.

Table 2. Experimental effective area *A* and its standard uncertainty *u*(*A*) of FPG at nominal pressures *p* in gauge and absolute modes.

|  |  |  |  |  |
| --- | --- | --- | --- | --- |
| p / kPa | Gauge mode | | Absolute mode | |
| *A* / cm2 | *u*(*A*) / cm2 | *A* / cm2 | *u*(*A*) / cm2 |
| 2 |  |  | 9.80592 | 4.8⋅10-4 |
| 3 | 9.80624 | 2.6⋅10-4 | 9.80601 | 2.5⋅10-4 |
| 5 | 9.80620 | 1.6⋅10-4 | 9.80600 | 1.8⋅10-4 |
| 6 | 9.80611 | 8.8⋅10-5 | 9.80602 | 1.8⋅10-4 |
| 8 | 9.80607 | 1.1⋅10-4 | 9.80601 | 1.5⋅10-4 |
| 10 | 9.80614 | 8.3⋅10-5 | 9.80601 | 9.4⋅10-5 |
| 11 | 9.80617 | 7.5⋅10-5 | 9.80606 | 1.0⋅10-4 |
| 13 | 9.80613 | 7.2⋅10-5 | 9.80598 | 9.4⋅10-5 |
| 15 | 9.80610 | 8.0⋅10-5 | 9.80599 | 8.7⋅10-5 |

In Figure 7, these experimental results are shown together with the results of a theoretical determination of FPG’s effective area based on the dimensional properties of the PCA, calculated using rarefied gas dynamics (RGD) methods and with a relative standard uncertainty of 4⋅10-6. Figure 7 also shows the effective area reported by the manufacturer with a relative standard uncertainty of 13⋅10-6. All experimental and theoretical effective areas agree well in that the experimental standard uncertainties are lower than typical FPG uncertainties [30].

The results of the pressure comparison between the FPG and the HgM are presented for gauge mode and absolute mode in Figures 8 and 9, respectively.

The pressure differences measured in gauge mode have a more pronounced scatter than in absolute mode since the gauge-pressure measurements were affected by the instability of the atmospheric pressure. Because the time constants of the FPG and the HgM differed considerably, the temporal reactions of the two instruments to the ambient pressure variations were significantly different, leading to the pressure differences observed in the experiment. In addition, the large volume of the tubes connecting the FPG to the HgM made it difficult to control the pressure during the measurements. Despite this, almost all pressure differences are covered by the expanded uncertainty ( = 2) of the HgM.

Figure 7. Experimental (Exper.), RGD-calculated (Theor.) and manufacturer’s (Manuf.) effective areas of the FPG in absolute (abs.) and gauge (gauge) modes.

Figure 9. Differences in absolute pressures measured with FPG and HgM.

Figure 8. Differences in gauge pressures measured with FPG and HgM.

On the contrary, in absolute-pressure mode, the measurements were free of the ambient pressure effect, and the results were characterised by a much lower scatter of the pressure differences (see Figure 9). The differences in the pressures measured with the FPG and those measured with the HgM are much smaller than the expanded uncertainty of the HgM [26], which demonstrates the consistency of the two pressure standards.

In the lowest pressure range of the FPG operation (from 3 Pa to 300 Pa), consistency was verified via a comparison with the SES in the absolute-pressure mode. The results of the comparison of the two pressure standards in terms of the pressure differences and their uncertainties are shown in Figure 10.

Figure 10. Differences in absolute pressures measured with FPG and SES.

Above 30 Pa, the pressure differences agree with the standard uncertainty ( = 1) of the SES. At lower pressures, the scatter of the pressure differences is greater than the standard uncertainty of the SES, but it is in agreement with the expanded uncertainty ( = 2) and allows conclusions to be drawn about the consistency of the two standards at 10 mPa.

1. Conclusions

For the theoretical and experimental effective area of the PCA, as well as for pressures generated by the FPG and the three reference standards based on different operation principles (the pressure balance, the mercury manometer and the static expansion system), all the results demonstrate full agreement within the expanded uncertainties ( = 2) of the three standards. Thus, the FPG, which has already been characterised as a primary pressure standard, is validated experimentally. For the pressure range from 3 Pa to 15 kPa, the FPG standard uncertainty is presently estimated as 10 mPa + 6⋅10‑6 × .

Acknowledgement

The first author wishes to thank the Egyptian study mission for their support and the funding of his scientific mission at PTB.

The research presented in this article was carried out within EMPIR, which is jointly funded by the EMPIR-participating countries within EURAMET and the European Union.

References

1. A. Ooiwa, Development of a highly stable air piston pressure gauge with non-rotational piston, Flucom'85, Tokyo, Japan, 2 September 1985, pp. 959-964.
2. A. Ooiwa, Novel nonrotational piston gauge with weight balance mechanism for the measurement of small differential pressures, Metrologia 30 (1993/94) pp. 607-610.
3. P. Delajoud, M. Girard, A force balanced piston gauge for very low gauge and absolute pressure, Proc. of the 10th Intern. Metrology Congress, St Louis, France, 22 – 25 October 2001.
4. R. Haines, M. Bair, Application of a new method for the automated calibration of very low gauge and absolute pressures in a commercial calibration laboratory, Meas. Sci. Conf., Anaheim CA, USA, January 2002.
5. A. Hashad, S. Ehlers, O. Jusko, W. Sabuga, Characterization of a force-balanced piston gauge as a primary pressure standard, Measurement 131 (2019) pp. 723-729.
6. S. Naris, N. Vasileiadis, D. Valougeorgis, A. Hashad, W. Sabuga, Computation of the effective area and associated uncertainties of non-rotating piston gauges FPG and FRS, Metrologia 56 (2019) pp. 1-10.
7. F. M. Sharipov, V. D. Seleznev, Rarefied gas flow through a long tube at any pressure ratio, J. Vac. Sci. & Technol. A 12 (1994) pp. 2933-2935.
8. V. A. Titarev, E. M. Shakhov, Rarefied gas flow through a long circular pipe into vacuum, AIP Conf. Proc. 1501 (2012) pp. 465‑472.
9. F. Sharipov, Rarefied Gas Dynamics. Fundamentals for Research and Practice, Wiley-VCH, Berlin, 2016, ISBN 978-3-527-41326-3.
10. P. Otal, J. C. Legras, Metrological characterization of a new standard, in absolute and gauge pressure modes, in the range 1 Pa to 15 000 Pa, Metrologia 42 (2005) pp. 216-219.
11. M. Rantanen, S. Semenoja, Results on the effective area of the FPG-type digital piston manometer of MIKES, Metrologia 42 (2005) pp. 165-168.
12. M. Rantanen, S. Semenoja, G. Peterson, J. Busk, Low gauge pressure comparisons between MIKES, Metrosert and FORCE Technology, Report on EUROMET Project 921, MIKES Publication J2/2007.
13. D. A. Vijayakumar, O. Prakash, R. K. Sharma, Establishment of a force balanced piston gauge for very low gauge and absolute pressure measurements at NPL, India, J. Phys. Conf. Ser. 390 (2012) 012013.
14. R. Haines, M. Bair, A method of traceability for a FPG8601 force balanced piston gauge to define pressures in the range from 1 Pa to 15 kPa in gauge and absolute measurement modes, Proc. of the XIX IMEKO World Congress, Lisbon, Portugal, 6 – 11 September 2009, pp. 2071-2076.
15. J. H. Hendricks, D. A. Olson, 1-15,000 Pa Absolute mode comparisons between the NIST ultrasonic interferometer manometers and non-rotating force-balanced piston gauges, Measurement 43 (2010) pp. 664-674.
16. J. H. Hendricks, D. A. Olson, NIST experience with non-rotating force-balanced piston gauges for low pressure metrology, Proc. of the XIX IMEKO World Congress, Lisbon, Portugal, 6 – 11 September 2009.
17. Z. Krajíček, M. Bergoglio, K. Jousten, P. Otal, W. Sabuga, S. Saxholm, D. Pražák, M. Vičar, Final report on EURAMET.M.P-K4.2010: Key and supplementary comparison of national pressure standards in the range 1 Pa to 15 kPa of absolute and gauge pressure, Metrologia 51 (2014) Tech. Suppl. 07002.
18. J. Tesar, P. Repa, D. Prazak, Z. Krajicek, L. Peksa, The new method of traceability of a force-balanced piston gauge used as primary vacuum standard, Vacuum 76 (2004) pp. 491-499.
19. J.-S. Chang, S.-G. Lee, Uncertainty evaluation of low pressure by means of cross-float, Proc. of the XX IMEKO World Congress, Busan, Rep. of Korea, 9 – 14 September 2012.
20. W. Sabuga, T. Priruenrom, R. Haines, M. Bair, Design and evaluation of pressure balances with 1⋅10-6 uncertainty for the Boltzmann constant project, PTB-Mitteilungen 121 (2011) pp. 256-259.
21. W. Sabuga, Pressure measurements in gas media up to 7.5 MPa for the Boltzmann constant redetermination, PTB-Mitteilungen 121 (2011) pp. 247-255.
22. T. Zandt, W. Sabuga, C. Gaiser, B. Fellmuth, Measurement of pressures up to 7 MPa applying pressure balances for dielectric-constant gas thermometry, Metrologia 52 (2015) pp. S305-S313.
23. Non-automatic weighing instruments, Part 1: Metrological and technical requirements – Tests, International Recommendation OIML R 76-1, Edition 2006 (E).
24. FPG8601™/VLPC™, Operation and maintenance manual, DH Instruments, a Fluke company, 2007.
25. H. Bauer, Die Darstellung der Druckskala im Bereich von 0.01 bar bis 2 bar durch das Quecksilber-Normalmanometer der PTB, PTB-Mitteilungen 89 (1979) pp. 248-255.
26. J. Jäger, Use of a precision mercury manometer with capacitance sensing of the menisci, Metrologia 30 (1993/94) pp. 553-558.
27. M. Perkin, A. Picard, M. Lecollinet, K. Fen, M. Sardi, A. Miiller, A. Agarwal, M. Jescheck, C. Wüthrich, Final report on CCM key comparison CCM.P-K2: pressure (10 kPa to 120 kPa) absolute mode, Metrologia 45 (2008) Tech. Suppl. 07002.
28. T. Bock, H. Ahrendt, K. Jousten, Reduction of the uncertainty of the PTB vacuum pressure scale by a new large area non-rotating piston gauge, Metrologia 46 (2009) pp. 389-396.
29. W. Steckelmacher, The calibration of vacuum gauges, Vacuum 37 (1987) pp. 651-657.
30. M. Bair, P. Delajoud, Typical pressure measurement uncertainty defined by an FPG8601 force balanced piston gauge, DH Instruments, a Fluke company, Technical Note 2090TN05, Rev. 26 March 2004.