Assessing the level of confidence for expressing extended uncertainty through control errors on the example of a model of a means of measuring ion activity

Oleksandr Vasilevskyi1

1 Vinnytsia National Technical University, 95 Khmelnitsky Shose str., 21021, Vinnytsia, Ukraine

Section: RESEARCH PAPER

Keywords: measurements; uncertainty; level of confidence; control errors; activity of ions; expanded uncertainty

Citation: Thomas Bruns, Dirk Röske, Paul P.L. Regtien, Francisco Alegria, Template for an IMEKO event paper, Acta IMEKO, vol. A, no. B, article N, month year, identifier: IMEKO-ACTA-0A (year)-0B-0N

Section Editor: name, affiliation

Received month day, year; In final form month day, year; Published month year

Copyright: © year IMEKO. This is an open-access article distributed under the terms of the Creative Commons Attribution 3.0 License, which permits unrestricted use, distribution, and reproduction in any medium, provided the original author and source are credited

Funding: This work was supported by Measurement Science Consultancy, The Netherlands

Corresponding author: Oleksandr Vasilevskyi, e-mail: o.vasilevskyi@gmail.com

Introduction

In view of the above, there is a need to develop methods for estimating the uncertainty of dynamic measurements that would meet international requirements for estimating the characteristics of the quality of measurements, which is a topical scientific task in the field of metrology International standards JCGM 100:2008 (GUM 1995 with minor corrections), JCGM 102:2011, JCGM 104:2009, ISO/IEC Guide 98-1:2009 «Uncertainty of measurement – Part 1: Introduction to the expression of uncertainty in measurement», support the use of combined uncertainty uc (y) as a parameter for quantifying the uncertainty of the measurement result [1 - 4].

Although the combined uncertainty uc (u) can be widely used to express the uncertainty of the measurement result, however, in some cases: in trade, industry and regulatory acts, and when it comes to health and safety, it is advisable to further specify an interval measure of uncertainty that determines the interval for the measurement result. An additional measure of uncertainty, which corresponds to interval estimation of uncertainty, is called expanded uncertainty.

Therefore, the task of evaluating the trust level for calculating the expanded measurement uncertainty with the use of a particular measurement tool is relevant.

The purpose of the article is to develop a mathematical apparatus for estimating the confidence level in calculating the expanded uncertainty of measuring the activity of ions taking into account the manufacturer's (developer's) and consumer control errors, which will enable to establish the interval around the measurement result of the ion activity, within which the majority of the distribution of values which can be attributed to the measured value.

MAIN MATERIALS OF THE RESEARCH

In literary sources [1 - 9], only a partial consideration is given to ways of establishing a trust level for calculating the expanded uncertainty of measurement. The mathematical apparatus that would allow reasonable confidence level set in the concept of measurement uncertainty is not described. It is therefore advisable to propose and describe a methodology for estimating trust level based on the control errors of the manufacturer and the consumer, which will allow to establish the value of the coefficient of coverage k for the calculation of the expanded uncertainty of measurement on an example of the use of the developed means for measuring the activity of ions.

The confidence level for calculating the expanded uncertainty is proposed based on the control errors of the manufacturer and the consumer by the formula

,	                                                   (1)

where α is control errors of the manufacturer; β is control errors of the consumer; Pn is the total value of control errors.

The procedure for determining the confidence level is described on the basis of using the combined uncertainty of the results of reusable measurements of ion activity using a measuring instrument built on the principle of converting voltage to frequency [10, 11], which is described by such a transformation equation

,                 (2)

where рХі is activity of ions;  is temperature coefficient of steepness, equal to 198,4·10-3; t is the temperature of the medium under investigation; k2  is gain scaling factor (DA1, DA2, fig. 1); Upow is value of the reference voltage of the voltage converter in the frequency (10 V); τ = RC is the time constant of the voltage converter in a frequency that is used to set the full scale output frequency range (R=1 kOm, С=47 µF); f0 is frequency of the quartz resonator microcontroller (20 МHz)  is the standard potential of the reference electrode at the initial isopotential point.;  is charge of the i-th ion [11].

The electrical functional circuit of the ion activity measuring device built on the conversion of voltage to frequency is shown in Fig. 1.

The general law of uncertainty of measuring ion activity depends on many factors such as activity interfering ions limited properties of ion-selective electrodes, the presence of measurement error in temperature, zero drift, instability of power supply, etc., among which is difficult to identify dominant [10]. This allows us to adopt the law of distribution of the centered value of the error of measuring the activity of ions in the normal, which we describe with the expression

,         (3)

where is estimation of ion activity value; uc(pX) is combined uncertainty of measurement ion activity.

The value of the combined uncertainty of measuring ion activity when using a measuring device built on the basis of the conversion of voltage to frequency is calculated by the formula

,                                             (4)

,                             (5)

where uA(pX) is evaluation of standard uncertainty type A; ui(pX) is components of the measurement uncertainty of activity of ions of type B with the respective sensitivity coefficients сi; uj(t) is components of temperature measurement uncertainty of type B with corresponding sensitivity coefficients cj; r(pX, t) is the correlation coefficient between ion activity (pX) and temperature (t) [10-14].

Taking into account the transformation equations (2), the combined uncertainty of type B is determined by the formula

,                                                         (6)

,                             (7)

where  is the equation of transformation of a measuring channel of temperature (Fig. 1) [10];  Es=ІR0 is sensor supply voltage (R0 is sensor resistance at temperature 0 °С); k1  is gain scaling factor (DA3, fig. 1);

[pX-1] is sensitivity factor for ion activity at a temperature of 25 °С;             [V-1] is coefficient of voltage sensitivity of the standard potential of the reference electrode;              [°С-1] is coefficient of sensitivity at an additional measured temperature;               [V-1] is coefficient of voltage sensitivity of the source of stable power supply of the voltage converter to the frequency;  is the number of pairwise measurements of the temperature and ion activity;  [pX] is uncertainty of type B, which is due to the presence of a limited class of accuracy (=0,7%) of the primary electrode;                       [mV] is uncertainty of type B, which is caused by the instability (=1,65 mV) of the standard potential of the reference electrode;                    [°С] is uncertainty of type B, which is caused by the existence of a bias voltage (= 0,2·10-3) of the operational amplifier when the temperature deviates ( = 5 °С) from the nominal value;  [mV] is uncertainty of type B, which is caused by instability of the power supply source of the reference voltage ( = 2mV);   [pX] is uncertainty of type B, which is caused by the nonlinearity () of the voltage converter in a frequency (AD650) [10 - 14].

The values of the combined uncertainty are calculated in [10, 14] and its maximum value for measuring the activity of phosphate ions is 17,21·10-3 pX in the range of measurements from 6 to 0.3 pX.

A compatible two-dimensional confidence level density when measuring the activity of ions, taking into account the allowable deviation of the measurement error ε, which is established by the consumer, is described by the expression

.   (8)

In most practical cases, as noted in [9, 15], the admissible deviation ε of the parameter which is controlled by the user, depending on the combined uncertainty of the measurement results uc(y), is determined by the formula

,	                    (9)

where d is width of the tolerance field determined by the values of the upper Xmax and lower Xmin range of measurement of the physical quantity.

Substituting the values of the maximum combined uncertainty of measuring the ion activity uc(pX)=17,21·10-3 pX and the value of the upper Xmax = 6 pX and the lower Xmin= 0,3 pX of the measurement limit in equation (9), we obtain the values of the permissible error set by the consumer, which is ε = 1.81%.

The uncertainty uε of the tolerance field of the controlled parameter at which the result can be considered reliable in practice is recommended to be equal to 6uc(y) [13, 15-20].  Thus accept uε = 6·uc(y) = 6*17,21·10-3 = 103,26·10-3 pX.

Taking into account the expression (8), the control errors of the manufacturer α is estimated by the formula

,(10)

the control errors of the consumer β is estimated by the formula

,(11)

The admission field is the value of the measured value ΔpX, in this case, this is the activity of the pX ions, we define by the formula

, 		                                   (12)

where is еstimation of activity of ions on the upper measurement range; δmax is the maximum relative measurement error (for the measuring channel of ion activity this value is 0.7% in the range of measurements from 6 to 0.3 pX).

Consequently, to calculate the control errors of the manufacturer and the consumer, the field of admission ∆pX we will calculate by the formula (12), which when adding the corresponding numerical values is equal to 2.1·10-3 pX.

Substituting calculated tolerances in the formula for estimating control errors manufacturer (10) and consumer (11) and solving them using the Maple 12 math package we obtain the following numerical values: α = 0.24·10-67 (α≈0), β = 0.00196. The total value of the control errors is Pn = 0.00196, and the confidence level for calculating the expanded uncertainty of the measurement, according to formula (1) will be equal to D = 1-Pn = 1-0.00196 = 0.998.

Characteristics of the change of control errors of the manufacturer and the consumer depending on the parameter μ = uε/uc(pX), which establishes the relationship between the permissible uncertainty of uε, which is specified by the consumer (normative documents) and combined uncertainty of uc(pX), which is set by the manufacturer (developer) of measuring device for uε < uc(pX), uε≈uc(pX), uε > uc(pX) are shown in Fig. 2.

Fig. 1. The electrical functional circuit of the ion activity measuring device

Thus, based on the constructed characteristic of the change in reliability (Figure 2), which is obtained by calculating the control errors of the consumer and the manufacturer, it is calculated, within which the majority of the distribution of values that can be attributed to the measured value is likely to be located in depending on the value of the accepted tolerance for the monitored parameter.

As can be seen from Figure 2, with a tolerance of 6 combined uncertainties, the confidence level is 99.8%.

Given the level of confidence based on control errors, the expanded uncertainty can be calculated from the formula

,   	                 (13)

where kp is coverage ratio for the established confidence level (p=D), which is taken from the Student's table.

Thus, the coverage factor kp with the number of degrees of freedom 30 and the confidence probability of 99.8% is 3.385, and with the number of degrees of freedom 10 and the probability 99.8% - 4.14. Accordingly, substituting these values of the coverage factor in equation (13), the expanded uncertainty will be is U99,8=71,25·10-3 pX for the number of degrees of freedom 10 and U99,8=58,26·10-3 pX for the number of degrees of freedom 30.

Conclusions

The work describes the characteristics of the change in the control errors of the manufacturer and the consumer on the basis of which a general characteristic of the change in the control errors that occurs during the measurement is obtained the activity of ions on the basis of which it is possible to determine graphically the confidence level within which most of the distribution of values obtained by measuring the activity of ions is probably located.

The obtained nomograph allows one to determine with a high probability the confidence level at a given tolerance for the controlled parameter to determine the coverage factor in calculating the extended uncertainty of ion activity measurement.

The described approach for determining confidence probability on the basis of control errors can be applied for any types of measurements provided a separate calculation of the control errors of the manufacturer of the measuring instrument and the consumer (errors of the first and second kind).

Fig. 2. Characteristic of the change of control errors depending on the parameter μ

References

BIPM, IEC, IFCC, ISO, IUPAC, IUPAP and OIML «Evaluation of measurement data – Guide to the expression of uncertainty in measurement», Joint Committee for Guides in Metrology, Bureau International des Poids et Mesures, JCGM 100:2008, 2008, 120 р.

BIPM, IEC, IFCC, ISO, IUPAC, IUPAP and OIML "Evaluation of Measurement Data – Supplement 2 to the 'Guide to the Expre ssion of Uncertainty in Measurement' – Extension to any number of output quantities", Joint Committee for Guides in Metrology, Bureau International des Poids et Mesures, JCGM 102:2011, 2011.

BIPM, IEC, IFCC, ISO, IUPAC, IUPAP and OIML "Evaluation of measurement data – An introduction to the «Guide to the expression of uncertainty in measurement» and related documents", Joint Committee for Guides in Metrology, Bureau International des Poids et Mesures, JCGM 104:2009, 2009.

ISO/IEC Guide 98-1:2009, Uncertainty of measurement – Part 1: Introduction to the expression of uncertainty in measurement (ISO, Switzerland, 2009).

Vasilevskyi O.M. Calibration method to assess the accuracy of measurement devices using the theory of uncertainty, International Journal of Metrology and Quality Engineering, 2014, 5.04: 403.

IEC GUIDE 115-2007, Application of uncertainty of measurement to conformity assessment activities in the electrotechnical sector (IEC, Switzerland, 2007).

ISO/IEC 17025:2005, General requirements for the competence of testing and calibration laboratories (ISO, Switzerland, 2005).

Vasilevskyi O.M. Methods of determining the recalibration interval measurement tools based on the concept of uncertainty, Technical Electrodynamics, 6, 81 (2014).

Suelí Fischer  Beckert, Wagner Saucedo  Paim, Critical analysis of the acceptance criteria used in measurement systems evaluation, International Journal of Metrology and Quality Engineering, 8:23 (2017), DOI: 10.1051/ijmqe/2017016.

Vasilevskyi O.M., Didych V.M. Elements of the theory of constructing potentiometric means of measuring control of ion activity with an increased probability [monograph], Vinnytsa: VNTU, 2013, 176 p., available in:  http://vasilevskiy.vk.vntu.edu.ua/file/d23a8f23c4fdc081c30ebb09bc7182e6.pdf.

Didych V.M., Vasilevskyi O.M., Podzharenko V.O. Potentiometer facilities of ions activity measurement of humus elements in soil, Visnyk of Vinnytsia Politechnical Institute, № 5, 2008, pp. 5-10, available in: https://visnyk.vntu.edu.ua/index.php/visnyk/article/view/639/638.

O. M. Vasilevskyi, Metrological characteristics of the torque measurement of electric motors, International Journal of Metrology and Quality Engineering, 8, 7 (2017).

Oleksandr Vasilevskyi, Volodymyr Didych, Anna Kravchenko, Maksym Yakovlev, Iryna Andrikevych, Dmytro Kompanets, Yevhen Danylyuk, Waldemar Wójcik, Askhat Nurmakhambetov, Method of evaluating the level of confidence based on metrological risks for determining the coverage factor in the concept of uncertainty, Proceedings Volume 10808, Photonics Applications in Astronomy, Communications, Industry, and High-Energy Physics Experiments 2018, 108082C, https://doi.org/10.1117/12.2501576.

Vasilevskyi O.M., Kucheruk V.Y., Bogachuk V.V., Gromaszek K., Wójcik W., SmailovaS., Askarova N. The method of translation additive and multiplicative error in the instrumental component of the measurement uncertainty, Proc. SPIE 10031, Photonics Applications in Astronomy, Communications, Industry, and High-Energy Physics Experiments, 2016, 1003127 (September 28, 2016), DOI:10.1117/12.2249195.

Wheeler D.J., An Honest Gauge R&R Study. Manuscript No. 189, 2009, available in: http://www.spcpress.com/pdf/DJW189.pdf.

Vasilevskyi, O.M., Kulakov, P.I., Dudatiev, I.A., Didych, V.M., Kotyra Andrzej, Suleimenov Batyrbek, Assembay Azat, Ainur Kozbekova Ainur, Vibration diagnostic system for evaluation of state interconnected electrical motors mechanical parameters, Proc. SPIE 10445, Photonics Applications in Astronomy, Communications, Industry, and High Energy Physics Experiments 2017, 104456C (August 7, 2017); doi:10.1117/12.2280993.

Vasilevskyi, O.M., Kucheruk, V.Y., Bogachuk, V.V., Gromaszek, K., Wójcik, W., Smailova, S., Askarova, N., The method of translation additive and multiplicative error in the instrumental component of the measurement uncertainty, Proc. SPIE 10031, Photonics Applications in Astronomy, Communications, Industry, and High-Energy Physics Experiments, 2016, 1003127 (September 28, 2016), DOI:10.1117/12.2249195.

Wheeler D.J., An Honest Gauge R&R Study. Manuscript No. 189, 2009, available in: http://www.spcpress.com/pdf/DJW189.pdf.

Podzharenko, V.O., Vasilevskyi, O.M., Diagnostics of technical condition of electromechanical systems for the logarithmic decrement, Proceedings of Donetsk National Technical University (2005).

Suelí Fischer  Beckert, Wagner Saucedo  Paim, Critical analysis of the acceptance criteria used in measurement systems evaluation, International Journal of Metrology and Quality Engineering, 8:23 (2017).

# Text boxes

ABSTRACT

A method for estimating the level of confidence for determining the coverage factor based on control errors is proposed using the example of using information on tolerances and uncertainty of measuring the activity of ions, which allows to establish a reasonable interval around the measurement result, within which most of the values ​​that can be justified are assigned to the measured value..